## Slide 1: HZB sample test cavity for investigating SRF materials
RaSTA
Rapid Sample Test Apparatus
IFAST/ISAS Meeting
Berlin, 02 – 03 December 2024
Felix Kramer

## Slide 2: RaSTA: Pre-testing in less time
Quadrupole Resonator (QPR)
  High quality measurements, but ≥ 2 weeks turnaround time
Turnaround time mostly set by infrastructure
  Cleanroom work for sample preparation, pre-assembly and mounting
  Vertical LHe cryostat, radiation protection measures
Wishlist for a new sample test cavity
  Use QPR samples (comparability, calorimetric measurement)
  Direct measurement of RBCS and Rres
  ≥ 1 sample per week
  Operation without radiation protection measures
    Low RF field level only
  Fit in “small” cryostat
    Compact cavity
    Increased frequency (compared to QPR)
RastA
 RaSTA: Rapid Sample Test Apparatus

## Slide 3: RaSTA: RF parameters
Pillbox-like cavity design
RF mode: TM020 at 4.8 GHz
H-field: zero-crossing at coaxial gap
Direct measurement of Rres possible
Operation without radiation protection  Limit: 5 kV
  Bsample = 8.5 mT
  Psample = 0.4 W
  Ptotal = 35 W
German patent 10 2021 123 261
RaSTA
Temperature | Nb RBCS | Nb3Sn RBCS
4.2 K | 7300 nΩ | ≈ 100 nΩ
1.8 K | 62 nΩ | << 1 nΩ
1.5 K | 10 nΩ | << 1 nΩ

## Slide 4: RaSTA: RF parameters
Pillbox-like cavity design
RF mode: TM020 at 4.8 GHz
H-field: zero-crossing at coaxial gap
Direct measurement of Rres possible
Operation without rad. protection  Limit: 5 kV
  Bsample = 8.5 mT
  Psample = 0.4 W
  Ptotal = 35 W
German patent 10 2021 123 261
RaSTA
H
Temperature | Nb RBCS | Nb3Sn RBCS
4.2 K | 7300 nΩ | ≈ 100 nΩ
1.8 K | 62 nΩ | << 1 nΩ
1.5 K | 10 nΩ | << 1 nΩ

## Slide 5: RaSTA: First prototype
In-house modification of standard components
Niobium coating at University of Siegen
SRF cavity with copper gaskets
  Thermal decoupling of cavity and sample
RaSTA 1.0

## Slide 6: First results from a niobium sample
RF spectrum looks good, modes are well separated
Calorimetric measurement successful
Open issues
  Adjust antennas and coupling
  Identify systematic error sources
  Deal with high surface resistance (parasitic losses?)
Small cavity  prototyping is possible !
RaSTA 1.0

## Slide 7: RaSTA 2.0
Rasta 2.0
Improved cooling of RF surfaces
  Copper as base material for all parts with RF fields
  Reduced material thickness
  Higher RF fields possible
  Thermally more stable
  Proper coupling measurements
Dummy RF ports
  Less higher mode components, less parasitic losses
Reduced cold mass (= stainless steel)
  Cavity: Only two ring-shaped flanges
  No sample adapter flange
  Less LHe needed for cooldown
Redesign of sample attachment
  No sample adapter flange, direct mounting to host cavity
  No cleanroom pre-assembly  improved turnaround time
  Longer coaxial line  less parasitic losses

## Slide 8: RaSTA
Based on QPR experience
Complementary tool, enables interative testing
Promising prototype results  optimized version in production
Summary

## Slide 9: Prototype cavity performance
Cooldown: Monitoring of cavity resonance
  Quality factor Q0 vs. temperature
  Separate sc transitions visible, cavity turns sc first
  Estimate RRR of niobium coating
Rasta 1.0
Time
Cavity nc
Cavity turns sc

## Slide 10: Prototype cavity performance
Cooldown: Monitoring of cavity resonance
  Quality factor Q0 vs. temperature
  Separate sc transitions visible, cavity turns sc first
  Estimate RRR of niobium coating
Rasta 1.0
Time
Cavity nc
Cavity turns sc
Sample turns sc

## Slide 11: Prototype cavity performance
Cooldown: Monitoring of cavity resonance
  Quality factor Q0 vs. temperature
  Separate sc transitions visible, cavity turns sc first
  Estimate RRR of niobium coating
Bulky stainless steel parts limit heat transport
  CW quench limit: ≈ 5 mT
Rasta 1.0
Time [s]
Sample Temperature [K]
RF Off
Quench

## Slide 12: Reflected Pulse Shapes
Current Test
B_peak = 0.5 mT
Clearly over coupled

## Slide 13: Reflected Pulse Shapes
Current Test
B_peak = 2.5mT
Under coupled?
Dip in power from zero crossing of voltage
But second peak is lower than first peak
Hypothesis: local thermal runaway on niobium film