## Slide 1: Unmixing Spectrum Analysis for Noble Gas Detection Charles-Philippe Mano - CEA/DAM/DIF charlesphilippe.mano@cea.fr
Thesis supervisors : - Frédérick Carrel (CEA/DRT)
- Hugues Paradis (CEA/DAM/DIF) - Luc Patryl (CEA/DAM/DIF)

## Slide 2: CTBT Context
The Comprehensive Nuclear-Test-Ban Treaty (CTBT) :
    International Monitoring System (IMS)
      Seismic
      Acoustic
      Hydroacoustic
      Radiological (particulates fission product : 80 stations gazeous fission products : 40 stations)
Stations Map for Radioxenons Measurement
ARIX ; SAUNA ; SPALAX ; Not yet installed stations (white).

## Slide 3
Even when the nuclear test is underground, some radioactive fission products can escape.
4 fission products of interest : 131mXe , 133mXe, 133Xe, 135Xe.
Radionuclide rejections during a nuclear test
Migration
Detection
Cavity
(1 kt)
Rejection
Atmospheric
Transport
Thousands of km
Hundreds of m
133Xe  1016 Bq, 135Xe  1017 Bq
1013 Bq, 1014 Bq
~10-2 Bq/m3
~10-3 Bq/m3
12000 m3/kt water steam
3000 m3/kt gaz (H2, CO2, CO…Xe)
 50 g/kt fission products (1019 Bq/kt)
133Xe , 135Xe
dilution in the atmosphere
very low activity to measure

## Slide 4: The radioxenon measurements method
Radionuclides of interest : important fission products from a nuclear explosion
Coincidence spectrometry to improve sensitivity
  Simultaneous detection of two events
  4 physical properties :
    Decay : b- and internal transition
    Desexcitation g, electron conversion and X-rays
Drawing of a coincidence 2D matrix
Detection of conversion an electron (EB) in coincidence with a photon (EA)
Detection of a b in coincidence with a photon (EA’)
b and e- spectrum
- decay scheme
Internal conversion drawing
Improvement of Signal-to-Noise ratio
Specific fingerprints for each radionuclide

## Slide 5: Set-up at the laboratory to test a new method
Picture of the system
Coincidence matrix with the 4 radioxenons
Development of a laboratory system (based on low resolution technology for photon measurements – NaI)
System description
  Photons g and X : NaI(Tl)
  Electrons and b : Silicon
  Electronic : Pixie-Net  4 synchronized channels
  Data : 2D matrix

## Slide 6
Current analysis for matrix coincidences : Regions of Interest (ROI)
214Bi
214Pb
135Xe
131mXe 133mXe
133Xe
Energy ]
Energy
  Method for high resolution systems at NDC
  One ROI for each radioxenons
    4 radioxenons
    Radon daughters (214Bi + 214Pb)
  Subtraction of interferences of each radionuclide
Cnet : net count in the ROI (matrix 6x1)
Cbrut : gross count in the ROI (matrix 6x1)
B : background in the ROI (matrix 6x1)
R : interference matrix (matrix 6x6)

## Slide 7: The overlap problem
133Xe
131mXe
133mXe
Low activity + overlap  Complex to analyse
 Research of a new method
7 / 17

## Slide 8
Hypothesis : multi-γ spectrum  linear combination of individual spectra*
  Spectral unmixing : calculation of individual abundances in the spectrum
  Data library of reference spectra : reference spectra depend of the detector, the radionuclide and the environment
No ROI analysis, no peak search  global analysis of the spectrum
  background used as a reference spectrum and add in the data library.
New method « Spectral unmixing » (1/3)
* Paradis H. et al. Spectral unmixing applied to fast identification of gamma-emitting radionuclides using NaI(Tl) detectors, Applied Radiation and Isotopes, 158 (2020) 109068.
Energy (chanel)
Probability
137Cs
Background
40K
60Co
Global spectrum
Global spectrum = Σ individual spectra

## Slide 9
Abundance of a certain radionuclide in the measured sample
Radionuclide signature, i.e. corresponding spectrum of a single radionuclide in the sample
The problem can be seen as an inverse problem : « We need to find the causes from the effects »
Indice i |  | 
1 | Spectral signature of Mvp | Abundance of Mvp
2 | Spectral signature of 137Cs | Abundance of 137Cs
3 | Spectral signature of 40K | Abundance of 40K
4 | Spectral signature of 60Co | Abundance of 60Co
Indice i |  | 
1 | Spectral signature of Mvp | Abundance of Mvp
2 | Spectral signature of 137Cs | Abundance of 137Cs
3 | Spectral signature of 40K | Abundance of 40K
4 | Spectral signature of 60Co | Abundance of 60Co
Global spectrum
New method « Spectral unmixing » (2/3)
Probability
137Cs
Background
40K
60Co
Global spectrum
Global spectrum = Σ individual spectra

## Slide 10
We want to determine abundances of radionuclides behind the observed spectrum s
We want to find the ai with respect to :
New method « Spectral unmixing » (3/3)
We need an estimation of the
Likelihood function :
We can deduce from the likelihood an estimator than can be computed according to :
10 / 17
Background
137Cs
60Co
Example of a spectrum composed of background, 137Cs and 60Co

## Slide 11
Reference spectra for each radionuclide
214Bi
214Pb
135Xe
133Xe
133mXe
131mXe
Energy
Energy
Energy
Energy
Energy
Energy
Energy
Energy
Energy
Energy
Energy
Energy
+
Building the data library to apply the spetral unmixing spectra

## Slide 12
Simulation of a mix of radionuclides by Monte-Carlo simulations
1000 simulations  uncertainties : standard deviation
Good agreement between activities obtained with the reference method and with unmixing method
Uncertainties improved
Results (1/4) : accuracy
RN | A simulated (mBq) | AROI (mBq) | Aunmix (mBq)
214Bi | 40,0 | 40,7 ± 6,2 | 38,9 ± 4,9
214Pb | 40,0 | 41,2 ± 5,7 | 39,3 ± 3,4
135Xe | 171,6 | 170,5 ± 3,8 | 169,9 ± 3,4
133Xe | 20,0 | 17,9 ± 13,5 | 21,0 ± 6,2
133mXe | 28,4 | 28,5 ± 3,4 | 28,0 ± 2,5
131mXe | 5,6 | 5,8 ± 3,4 | 5,6 ± 1,9
Example of a mix representative of a nuclear test
Activities estimated
Activity (mBq)
True values
ROI activites
Unmixing values

## Slide 13
Critical level (CL) and detection limit (DL) with risks
a and b equal to 5%
  Currie formula for ROI analysis :
  Monte-Carlo simulations for unmixing
Background analysis (300 counts in the matrix  8h measurement) :
Results (2/4) : sensitivity
RN | CLROI (mBq) | CLunmix (mBq)
135Xe | 1,2 | 0,8
133Xe | 10,4 | 0,2
133mXe | 1,6 | 0,3
131mXe | 1,6 | 0,4
RN | DLROI (mBq) | DLunmix (mBq)
135Xe | 2,7 | 1,8
133Xe | 23,0 | 1,1
133mXe | 4,6 | 1,3
131mXe | 4,6 | 1,4
CL
DL
Improvement of the sensitivity

## Slide 14
Concentration of Rn daughters
Results (3/4) : radon contamination
> 10
> 10
> 10
> 10
1000 2000 3000 4000
1000 2000 3000 4000
1000 2000 3000 4000
1000 2000 3000 4000
133Xe
133mXe
131mXe
CL with a equal to 5%
Variation of proportion of radon daughters
CL ↘ reliability
Critical level (counts)
135Xe

## Slide 15
Concentration of Rn daughters
Results (4/4) : radon contamination
DL with b = 5%
Variation of proportion of radon daughters
DL ↘  Sensitivity improved
DL less damaged by the presence of Pb/Bi-214  more robust if Rn is present in the sample.
1000 2000 3000 4000
1000 2000 3000 4000
1000 2000 3000 4000
1000 2000 3000 4000
135Xe
133Xe
133mXe
131mXe
Detection limits (counts)

## Slide 16
Adaptation of a new global method on two dimension spectra
  Use of a maximum of information
  Possibility to perform automatic analysis
Improvement of the sensitivity for radioxenon measurements
  CL improved of a factor of 10
  DL improved
  Method more robust with Rn daughter presence
          Improve the reliability of RXe detections in case of nuclear test
Conclusion
** Mano CP. et al. Algorithm development for low-level radioxenon 2D spectra analysis : a first case of study with spectral unmixing using a β-γ detector, Applied Radiation and Isotopes (2023)
**

## Slide 17
Comparison to be done with tools developed by CTBTO tools
Addition of sparsity constraints to improve the false positive responses
   Variation of the data library to keep just RN present
Automatic analysis of matrices from low-resolution system
  Comparison with officials CTBTO tools
 Maximizing network utilization
Outlook
Thank you
Coincidence matrices for each radionuclide of interest from low resolution system
charlesphilippe.mano@cea.fr