## Slide 1
Project Progress Review (PPR) 2-2024
WP- Laser
Presented by: Fatematuj Johora
January 20, 2025

## Slide 2: Objective: Laser Installation
Model of Optical Transport Line
Optical room design
Objectives:
To design the optical room and optical transport line.
To check the Laser parameters before transportation.
Laser Installation, order, and measurements.

## Slide 3: PERLE Photocathode laser-SATSUMA
The overall system is composed of:
Satsuma Controller
Satsuma Laser head
Stretcher Module
Chiller
PERLE photocathode Laser: SATSUMA from Amplitude.
Laser received and installed in Dec 2023.
Amplitude intervention:
Changed Peltier and temperature controller on July 21, 2024.
Changed pump diode and added water cooling system to the software in February, 2025.

## Slide 4: PERLE Photocathode laser-Burst Mode Constraints
The oscillator repetition rate is 40MHz.
Burst mode possible for 10MHz & 20MHz of Pulse Picker frequency.
AOM frequency: 10Hz - 2MHz.
AOM width: 25ns - 2500ns. & AOM Delay= 0ns - 2300ns.

## Slide 5: Beam Radius Measurements
79 cm
115 cm
235 cm
581 cm
Beam radius measured at different distances from the laser output @40MHz.
Total divergence is 0.8mrad.
Figure: Setup for Beam radius measurement
Figure : Beam profile at different distances (left); Beam radius vs Distance curve (right).

## Slide 6: Power Measurements Report -19 Feb 2025
The output power in SHG was measured for different %Transmission of AOM.
Figure: Laser power vs % Transmission AOM @10MHz (left); @40MHz (right). Blue data shows the actual power after SHG and the stretcher module and red shows the power from the software window.
Need to fix the problem in software maximum limit, asked to the Amplitude person.
After
intervention
Before
intervention

## Slide 7
100 kHz
10 kHz
10 kHz, 2 pulses
1 kHz
100 Hz
10 Hz
Measurements at low repetition rate- 17 February 2025
With Pulse Picker frequency 10MHz, different repetition rate for the external modulator was checked.

## Slide 8: Diagnostic box
A fixed aperture of 1mm.
We need 3-4mm.
Two boxes: Box1 will be placed before the photocathode and
Box 2 after the photocathode.
Box 1 was opened on 12 March 2025 and the aperture and
the zoom telescope were removed for tests.
Box1
Box2

## Slide 9: Diagnostic box- Measurement on 18 March 2025
Need to build a telescope for the optical room.
Using the zoom telescope different magnification m= 0.6, 1 & 1.4 tested for a distance of 13.3m from laser output.
Dx=1.7mm
After telescope @13.3m from the laser output
Dx=2.7mm
Dx=3.6mm
Before telescope
@1.36m from laser output
Dx=2.7mm

## Slide 10: Optical Room Preparations
Height of:
Laser output =982.3mm
Diagnostic box = 1400mm
Total optical path 13m
Designed the Optical room set-up and the optical transport line.
A hole is drilled on the optical room wall at a height of 1045mm. Initially proposed 995mm.

## Slide 11: Laser future and deadlines
Laser transport, installation and measurements required.
Laser has sensitive fibers, which we can’t remove. Transportation needs to be carefully done.

## Slide 12
THANK YOU