## Slide 1: Overcoming challenges in coincidence matrix analysis
H.-D. Lenouvel1, P. Gross1 , A. De-Vismes Ott1 , M. Morin1 , S. Topin1,2 , J. Aupiais1, H. Paradis1
1 CEA, DAM, DIF, F-91297 Arpajon, France
2 CNAM, Analyse Chimique et Bioanalyse, EPN-7, 292 rue Saint Martin, 75003 Paris, France
Corresponding author : hugues.paradis@cea.fr

## Slide 2: Introduction
1

## Slide 3
Comprehensive Nuclear-Test-Ban Treaty (CTBT)
321 stations and 16 laboratories:
4 types:
Seismic
Radionuclides
  80 particles
  40 noble gases
Hydroacoustic
Infrasound
Many nuclear test campaigns  Comprehensive Nuclear-Test-Ban Treaty (1996) [1]
  States :
    Signature : 187
    Ratification : 178
  High atmospheric dilution  Low radioactivity
  SPALAX
  SAUNA
IMS status (Dec 2023)
( : non-op. station)
XENON (1–4 samples/d)
  ARIX
Fig 1. Xenon stations Map of the IMS
Migration
Detection
Cavity
Atmospheric transport
Thousands of km
133Xe  1012 Bq
~10-3 Bq/m3
Fig 2. Example for an underground nuclear explosion of 1 kt eq. TNT [2]
 Radionuclide technology allows to qualify the nuclear nature of an event
[1] United Nations, The Comprehensive Nuclear-Test-Ban Treaty, Resolution CTBT/MSS/RES/1, 1996.
[2] Boxue Liu et al., Characterization of CTBT-Relevant Radioxenon Detections at IMS Stations Using Isotopic Activity Ratio Analysis, Pure Appl. Geophys. 180 (2023), 1521-1540

## Slide 4
Resources deployed by the CEA as part of the CTBT framework:
  Stations : SPALAX and SPALAX-NG [3]
  National Data Center (NDC)
  Certified Laboratory : FRL08 (Gamma3) [4]
Research for radionuclide measurements :
  Coincidence measurements (with the Gamma3)
  Data analysis
Radionuclide technology : context and stakes
Fig 3. SPALAX-NG
Fig 4. Gamma3 (left) & PIPSBox (right)
[3] S. Topin, et al. 6 months of radioxenon detection in western Europe with the SPALAX-New generation system - Part1: Metrological capabilities, J. Environ. Radioact. 225 (2020) 106442
[4] A. Cagniant, et al., An introduction to γ3 a new versatile ultralow background gamma spectrometer. Background description and analysis, Applied Radiation and Isotopes 98 (2015)
Notes: Desired characteristics :
  Nuclear weapons involves fission reactions  Detection of fission products and activation products

## Slide 5: Construction of the digital model
2

## Slide 6: Model creation with MCNP
 Use of sources of different sizes and compositions to provide a general calibration
Standard sources :
  Energy range covered : [22 - 2505 keV]
  Geometries :
  compressed or uncompressed particle filters (diameter of 110 mm)
  volumetric geometries (from 20 mL to 500 mL)
Modeling from :
  Technical drawings provided by the manufacturer
  Radiography of detectors
Fig 5. Radiography of a HPGe detector (BEGe5030P)
Fig 6. Display of a cross-sectional view of a detector modeled under MCNP.
Fig 7. Common sources geometry used by the laboratory
BP20
3M
RASA
PIPSBox geometry
Double Face

## Slide 7: Optimization of detector models
Chart: Energy (keV)
| 5.41 | 5.95 | 11.87 | 13.93 | 14.41 | 17.55 | 22.1 | 24.1 | 25.07 | 26.34 | 27.5 | 32.06 | 33.299999999999997 | 36.450000000000003 | 37.33 | 38.06 | 59.54 | 60 | 60 | 88.03 | 122.06 | 136.47 | 165.86 | 255.13 | 320.08 | 391.7 | 514 | 661.66 | 834.85 | 898.04 | 1115.55 | 1173.23 | 1332.49 | 1836.07 | 2734.09
Mean | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065 | 1.0065
Standard deviation (k=2) | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.83501 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445 | 0.95445
+2 sigma | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.1781 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586 | 1.0586
BP20 | 0 | 0 | 0 | 0 | 1.0152 | 0 | 0 | 0 | 0 | 0.96382 | 0 | 0 | 0 | 1.0307 | 1.029 | 0.99988 | 1.0044 | 1.0406 | 0.99959 | 0.9688 | 1.0045 | 1.0088 | 0.98948 | 0.97571 | 0.97131 | 0.9763 | 0.96841 | 0
3M | 0 | 0 | 0 | 0 | 0.88469 | 0 | 0 | 0 | 0 | 1.2171 | 0 | 0 | 0 | 0.97303 | 0.99843 | 1.0145 | 0.96188 | 1.0525 | 1.0134 | 1.0273 | 0 | 0.99888 | 1.0257 | 1.0211 | 0.99725 | 0.98106 | 0.99135 | 0.98879 | 0
Double_Face | 0 | 0 | 0 | 0 | 1.0405 | 0 | 0 | 0 | 0 | 1.0125 | 0 | 0 | 0 | 0 | 1.0199 | 1.0308 | 1.0266 | 1.0044 | 0 | 0.9998 | 1.0247 | 0 | 0.99948 | 0.98876 | 0 | 0 | 0 | 0 | 0 | 0
PIPSBox Geometry | 0 | 0 | 0 | 0 | 0.98805 | 0 | 0 | 0 | 0 | 1.0174 | 0 | 0 | 0 | 1.0335 | 1.0461 | 1.0244 | 0.97873 | 0.98177 | 1.0492 | 0.96621 | 0.96224 | 0 | 0
rasa | 0 | 0 | 0 | 0 | 1.2013 | 0 | 0 | 0 | 0 | 0.96434 | 0 | 0 | 0 | 0.96774 | 0.98676 | 0.98365 | 1.0071 | 1.0003 | 1.0889 | 1.0163 | 0 | 0.99852 | 1.0225 | 0.97426 | 0.99773 | 0.98954 | 0.98195 | 0.97477 | 0
1° Modeling of detectors
2° Optimization :
Mean ratio =
3° Validation of models
separation at 60 keV
Parameters to optimize:
Thickness of materials (dead layers, carbon window…)
Position of the source relative to the crystal detector
Crystal size
 Model validation enables simulation for calculating new detection efficiencies
Detection efficiency :
Validation of the digital model :
Fig 9. Detection efficiency ratio (EXP/MCNP) versus energy
Fig 8. Modeling of the configuration (ATHOS + ARAMIS) & PIPSBox
2 conditions must be met for the model:
0.99<Rmean<1.01
<20% below 60 keV
<10% above 60 keV
[5] A. Berlizov, A correlated particle source extension of a general purpose Monte Carlo N-Particle transport code, MCNP-CP Upgrade Patch Version 3.2, Institute for Nuclear Research National Academy of Sciences of Ukraine, 2012

## Slide 8: Radioxenon isotopes measurement
coincidences
3

## Slide 9: Gamma3 : configuration for the measurement of gaseous samples
Configuration :
  2 HPGe : detection of photons
  2 PIPS : detection of electrons
  The PIPSBoxTM is placed between ATHOS and ARAMIS detectors
coincidences :
Fig 10. (ATHOS+ARAMIS)&PIPSBox configuration
Fig 11. Depiction of the shape of the fingerprints
Fig 13. particle emission energy spectrum
Fig 12. Energy spectrum of the electron emitted by internal conversion

## Slide 10: Benefit of using coincidence detection
The benefit of using coincidences in the X-ray region:
Characteristic signatures of radionuclide (very low background)
Fig 15. -ray spectrum with a zoom on the X-ray region
Fig 14. Separation of the signatures of metastable radioxenon isotopes
Impossibility of using the X-rays due to peak overlaps

## Slide 11
Standard gaseous sample of radioxenon isotopes:
Validation for -ray spectrometry and for coincidence spectrometry
H.-D. Lenouvel, et al., Measurement of radioxenon isotopes for nuclear explosion detection using coincident detector calibrated by simulation, Applied Radiation and Isotopes. DOI : ARI 111886
RXe | (Bq) | k=2 (Bq) | (Bq) | k=2 (Bq) | Ecart (%) | score
131mXe | 94 | 10 | 84 | 13 | -10.6 | -0.60
133Xe | 204 | 23 | 205 | 19 | 0.8 | 0.06
133mXe | 10.2 | 1.1 | 9.1 | 1.4 | -11.3 | -0.63
135Xe | 217 | 22 | 226 | 21 | 4.0 | 0.29
RXe |  |  |  |  | Ecart (%) | 
131mXe | 94 | 10 | 84 | 13 | -10.6 | -0.60
133Xe | 204 | 23 | 205 | 19 | 0.8 | 0.06
133mXe | 10.2 | 1.1 | 9.1 | 1.4 | -11.3 | -0.63
135Xe | 217 | 22 | 226 | 21 | 4.0 | 0.29
Validation of the method on a gaseous sample
Measurement :
  (ATHOS+ARAMIS)&PIPSBox
  Duration : 57600 s
Analysis :
Efficiency calculation :
  Calculated with MCNP-CP
Radioxenon isotope | β/γ coincidence spectrometry
131mXe | 0.16 ± 0.02
133Xe | 0.20 ± 0.02
133mXe | 0.147 ± 0.023
135Xe | 0.176 ± 0.016
Fig 16. matrix obtained from the measurement
Tab 1. detection efficiencies for the 4 radioxenon isotopes
Tab 2. Comparison between reference activities and measured activities
Notes: *Validation with score according to the ISO 13528 standard
*Validation with 𝑬_𝒏 score according to the ISO 13528 standard
𝐸_𝑁=(𝐴_𝑖−𝐴_𝑟𝑒𝑓)/√(𝑈^2 (𝐴_𝑖 )+𝑈^2 (𝐴_𝑟𝑒𝑓 ) )

## Slide 12: Particulate fission products
4
coincidences

## Slide 13: Benefits of using coincidence detection (example with 134Cs)
Identification of 134Cs using coincidences:
Measurement of a fresh fission products sample from CTBTO:
  Classic -ray spectrometry : complex spectra (~300 peaks)
  Coincidence spectrometry : characteristic signatures
High SNR values make it easier to qualify and quantify certain radionuclides
Fig 17. Decay scheme of 134Cs
Fig 18. Aerosol filter placed between two HPGe detectors with expected coincidence matrix signatures
Fig 19. Classic 𝛾-ray with zoom around 605 keV
Fig 20. Coincidence matrix with a zoom on the 134Cs
: Signal-to-Noise Ratio
Notes: Ruthenium-103 -> 610
Antimoine-124 -> 602
Antimoine-127 -> 603

## Slide 14: Validation of the method on a IAEA sample
Radionuclides | (Bq) | k=2 (Bq) | (Bq) | k=2 (Bq) | Deviation (%) | score
110mAg | 7.51 | 0.09 | 7.36 | 0.70 | -2.0 | -0.21
139Ce | 7.51 | 0.09 | 6.7 | 1.1 | -11.0 | -0.78
134Cs | 7.56 | 0.07 | 7.75 | 0.79 | 2.5 | 0.24
152Eu | 7.55 | 0.07 | 8.0 | 0.8 | 5.7 | 0.55
235U | 0.0747 | 0.0009 | 0.062 | 0.027 | -16.2 | -0.46
Radionuclides |  |  |  |  | Deviation (%) | 
110mAg | 7.51 | 0.09 | 7.36 | 0.70 | -2.0 | -0.21
139Ce | 7.51 | 0.09 | 6.7 | 1.1 | -11.0 | -0.78
134Cs | 7.56 | 0.07 | 7.75 | 0.79 | 2.5 | 0.24
152Eu | 7.55 | 0.07 | 8.0 | 0.8 | 5.7 | 0.55
235U | 0.0747 | 0.0009 | 0.062 | 0.027 | -16.2 | -0.46
Analysis :
Detecting radionuclides (134Cs) with poor visibility in -ray spectrometry
Validation for -ray spectrometry and for coincidence spectrometry
Tab 3. Comparison between reference activities and measured activities
Fig 21. matrix obtained from the measurement with a zoom on the 134Cs
Standard IAEA swipe sample
Measurement :
  ATHOS&ARAMIS
  Duration : 316800 s
ROI
Bckg
Efficiency calculation :
  Calculated with MCNP-CP
Fig 22. Zoom on a region where 235U coincidences may occur

## Slide 15: Conclusions
5

## Slide 16: Conclusion and perspectives
  Metrological connection of the digital models of detectors
    Determination of detection efficiencies through simulation
  Validation of a new method using simulation to qualify and quantify samples with coincidence measurements
    Gaseous samples
    Particle filters
RXe | (Bq) | à k=2 (Bq) | (Bq) | à k=2 (Bq) | score *
131mXe | 94 | 10 | 84 | 13 | -0.60
133Xe | 204 | 23 | 205 | 19 | 0.06
133mXe | 10.2 | 1.1 | 9.1 | 1.4 | -0.63
135Xe | 217 | 22 | 226 | 21 | 0.29
RXe |  |  |  |  | 
131mXe | 94 | 10 | 84 | 13 | -0.60
133Xe | 204 | 23 | 205 | 19 | 0.06
133mXe | 10.2 | 1.1 | 9.1 | 1.4 | -0.63
135Xe | 217 | 22 | 226 | 21 | 0.29

## Slide 17: Outlook : AI tool
Goal : Classification for particulate samples ( concidences) in the CTBT context
  ~90 radionuclides of interest (fission products and activation products)
Database (~1M):
  Memory-intensive matrix (8k*8k bins)  matrix (2k*2k bins)
  Fission products :
    Nucleonica website [6] : distribution from fission of 235U (High Enriched Uranium) with fast neutrons of 14 MeV
    Time-dependent evolution of this mixture from t0 to 1 year (Bateman equation) : 100 iterations
Convolutional Neural Network (CNN) :
  Inputs : Matrices (interpreted as images)
  A CNN identifies characteristic features within an image/spectrum
  Full spectral information is used but a large number of pixels have zero intensity
Fig 23. Topaze High-Performance Computing (HPC) system (CEA)
[6] Nucleonica website, www.nucleonica.com
Thank you.
Notes: Simulation of coincidences of each radionuclide by MCNP-CP
  Reduction of input matrix size (~2000*2000) with resolution binning (1 bin = 1*FWHM)
Simulation of 𝛾/𝛾 coincidences of each radionuclide by MCNP-CP
  Reduction of input matrix size (~2000*2000) with resolution binning (1 bin = 1*FWHM)

## Slide 18: Annexes
3

## Slide 19: Gamma3 : detection device for radionuclides present at trace levels in the environment
Gamma3 (A. Cagniant, 2015):
  3 -ray spectrometers (BEGeTM 5030P)
  Optimised shielding :
    Passive (Materials : Pb, B-PE…)
    Active (detectors : plastic scintillators)
     Minimization of background
Multi-detector spectrometer:
  Increase detection efficiencies
  Multi-channel acquisition electronics coincidence measurements
  Characteristic signatures of radionuclides
  Ultra low background
Fig 5. Gamma3 (left) & PIPSBox (right)
Fig 5. coincidence matrix

## Slide 20: score calculation
Validation with score according to the ISO 13528 standard
If should be taken as an indicator of successful performance

## Slide 21: Activity calculation and uncertainties for radioxenon isotopes measurements
Background contribution :
  Radiological blank (6 days) : < 0,01 counts per bin (negligible)
  Interference matrix that examines the contribution of each radionuclide in each ROI
Activity uncertainty = uncertainties propagation :
  Uncertainty in the gross number of counts (Poisson statistics):
  The model uncertainty is applied to detection efficiency uncertainties
Dead time :
Detection efficiencies are calculated with MCNP-CP
Gas transfer yield
Decay correction factors :
  Decay between the reference time and the beginning of the measurement
  Decay during the measurement
Active time
Interferences calculation :
The number of counts a radionuclide generates in the ROI () of another radionuclide (noted as ) .

## Slide 22: Activity calculation and uncertainties for aerosol filters
ROI
Activity calculation :
ROI sizes :
  Confidence interval covering over 99% of the events
Uncertainty in the gross number of counts (Poisson statistics):
Bckg
Background estimation :
  The number of background counts is normalized to the background area and scaled ()
Drawing of an ROI on a coincidence matrix
Energy in ATHOS
Energy in ARAMIS

## Slide 23: Validation of the method on a IAEA sample
Radionuclides | Considered energies (keV) | (Bq) | at k=2 (Bq) | (Bq) | at k=2 (Bq) | Deviation (%) | score
110mAg | 1384.3 & 884.7 1384.3 & 657.8 | 7.51 | 0.09 | 7.36 | 0.70 | -2.0 | -0.21
139Ce | 38.9 & 165.9 | 7.51 | 0.09 | 6.7 | 1.1 | -11.0 | -0.78
134Cs | 604.7 & 795.9 | 7.56 | 0.07 | 7.75 | 0.79 | 2.5 | 0.24
152Eu | 778.9 & 344.3 40.1 & 1112.1 | 7.55 | 0.07 | 8.0 | 0.8 | 5.7 | 0.55
235U | 15.3 & 143.8 15.3 & 185.7 | 0.0747 | 0.0009 | 0.062 | 0.027 | -16.2 | -0.46
Radionuclides | Considered energies (keV) |  |  |  |  | Deviation (%) | 
110mAg | 1384.3 & 884.7 1384.3 & 657.8 | 7.51 | 0.09 | 7.36 | 0.70 | -2.0 | -0.21
139Ce | 38.9 & 165.9 | 7.51 | 0.09 | 6.7 | 1.1 | -11.0 | -0.78
134Cs | 604.7 & 795.9 | 7.56 | 0.07 | 7.75 | 0.79 | 2.5 | 0.24
152Eu | 778.9 & 344.3 40.1 & 1112.1 | 7.55 | 0.07 | 8.0 | 0.8 | 5.7 | 0.55
235U | 15.3 & 143.8 15.3 & 185.7 | 0.0747 | 0.0009 | 0.062 | 0.027 | -16.2 | -0.46
Analysis :
Detecting radionuclides (Cs-134) with poor visibility in -ray spectrometry
Validation for -ray spectrometry and for coincidence spectrometry
Tab 3. Comparison between reference activities and measured activities
Fig 21. matrix obtained from the measurement with a zoom on the Cs-134
Low-level radiological swipe performed to detect fission and activation products, and assess nuclear material lineage :
Measurement :
  ATHOS&ARAMIS
  Time : 316800 s
ROI
Bckg
Efficiency calculation :
  Calculated with MCNP-CP

## Slide 24: Detection limits
Radionucléides | Spectrométrie : LD (mBq) | Spectrométrie : LD (mBq)
131mXe | 73 | 0.3
133Xe | 3.3 | 1.0
133mXe | 33 | 0.2
135Xe | 18.3 | 4.3
Radionucléides |  | 
131mXe | 73 | 0.3
133Xe | 3.3 | 1.0
133mXe | 33 | 0.2
135Xe | 18.3 | 4.3
Radionucléides | Spectrométrie : SD / LD (mBq) | Spectrométrie en : SD / LD (mBq)
108mAg | 2.03 / 4.2 | 0.39 / 2.9
110mAg | 2.44 / 5.0 | 3.67 / 26.4
60Co | 1.81 / 3.8 | 0.60 / 4.3
125Sb | 5.31 / 10.9 | 5.39 / 19.0
134Cs | 1.75 / 3.6 | 0.71 / 3.2
Radionucléides |  | 
108mAg | 2.03 / 4.2 | 0.39 / 2.9
110mAg | 2.44 / 5.0 | 3.67 / 26.4
60Co | 1.81 / 3.8 | 0.60 / 4.3
125Sb | 5.31 / 10.9 | 5.39 / 19.0
134Cs | 1.75 / 3.6 | 0.71 / 3.2
Calculation of decision thresholds (SD) and detection limits (LD) on a blank measurement :
  ISO 11929 standard
  Comparison between the classical spectrometry and the coincidence spectrometry
Particulate samples :
Noble gas samples :
For certain radionuclides :
very low 
 is particularly advantageous  a more sensitive qualification
spectrometry is particularly advantageous for metastable radioxenon isotopes
Tab 4. Comparison of SD and LD between classical spectrometry and spectrometry for particulate measurements
Tab 5. Comparison of SD and LD between classical spectrometry and spectrometry for particulate measurements