## Slide 1
LCLS-II-HE Upgrade Status / LCLS-II Operations Experience
James Maniscalco | SLAC National Accelerator Laboratory
TTC Saclay 2026
11 June 2026

## Slide 2: LCLS-II-HE Mission Need: Ultrafast High Power Hard X-rays
Deliver the ability to observe and understand the structural dynamics of complex matter at the atomic scale with hard x-rays, at ultrafast timescales, and in operational environments.
LCLS-II-HE provides high-rate, FEL radiation at ngstrom wavelengths
LCLS-II-HE
Mission Need Statement
LCLS-II-HE Mission Need Statement.pdf

## Slide 3: LCLS-II-HE Scope Overview
Add 23 additional cryomodules (L4 linac) to double the LCLS-II accelerator energy: 4 GeV to 8 GeV
Install new cryogenic distribution box and transfer line between the cryoplant and the new L4 linac
New long period soft X-ray undulator
Upgrade XPP instrument & enable upgrades of DXS, MFX, and CXI beamlines for MHz beam.
Construct new tunnel for future Low Emittance Injector (LEI)
Develop prototype SRF gun for the LEI

## Slide 4: Cryomodule Overview
  8x 1.3 GHz 9-cell cavities
  CW operation @ 20.8 MV/m, 2.0 K
  CM spec: 173 MV, 137 W

## Slide 5: Cryomodule Systems Summary Schedule

## Slide 6: Cavity Procurement
192 cavities procured from an industrial supplier 2021-2024
  900 ºC or 950 ºC heat treatment (depending on material qualification by single cell)
    Ensure optimal flux expulsion and avoid over-softening material
  2N0 doping recipe
  23 CMs × 8 cavities/CM = 184 cavities; additional 8 yield spares
Frequent & open communication with supplier was essential for success
  Weekly progress meetings
  Many in-person site visits
  Collaborative & collegial problem-solving approach
LCLS-II and pre-production cavity remediation/requalification effort
  11 cavities successfully recovered meeting HE cryomodule performance spec

## Slide 7: Cavity Procurement
Cavity qualification testing overseen by Cavity Technical Board (CTB)
  Weekly meetings to review cavity and cryomodule results and issues
Manufacture @ supplier
Analysis by cavity technical board
reporting
oversight
Vertical qual.
@ partner labs
Cryomodule
assembly
PASS
Analysis by cavity technical board
FAIL /
MARGINAL
minor issue: light rework at labs
major
issue
CTB: SRF experts from SLAC, JLab, and FNAL
Set aside for possible rework

## Slide 8: Vertical Test Performance – as of May 28, 2026
Average Q0 plotted to arbitrary 30 MV/m limit
  192 (all) production cavities tested, of which:
    183 qualify (95.3%)
      Including 10 “marginal” with Eacc > 21 MV/m
      Including 1 “marginal Q0” with Q0 > 2.2×1010
      125 qualified on first test (65%)
    8 disqualified (4.2%)
      3 with low Eacc
      1 with low Q0
      2 with persistent FE
      1 with HOM scratch
      1 with bellows damage
    1 awaiting re-test after high-pressure rinse (0.5%)
      1 with field emission
  Reprocessing to restore/improve gradient:
    HF Rinse / EP Rinse: 18/23 successful (78%)
    Supplier re-doping: 6/6 successful (100%)
  Recovery of 11 pre-existing cavities:
    2 vCM yield cavities
    9 remediated LCLS-II cavities
  10 qualified spares after conclusion of CM testing
Field-emitting cavities excluded

## Slide 9: HE Cavities vs. LCLS-II: Accelerating Gradient in VT
Average quench gradient: 22.0±3.9 MV/m
Average quench gradient: 27.7±3.1 MV/m
LCLS-II
HE vCM + production
>5 MV/m improvement
LCLS-II spec: 19 MV/m
LCLS-II-HE spec: 23 MV/m

## Slide 10: HE Cavities vs. LCLS-II: Quality Factor in VT
Average Q0:
3.18±0.44×1010
Average Q0:
3.25±0.38×1010
LCLS-II
HE vCM + production
Only considering cavities that reached at least 21 MV/m
Very few cavities below spec

## Slide 11: Quality Factor by Heat Lot
Material segregation strategy has been successful
  Only 6 out of 192 cavities (3.1%) with Q0(20.8 MV/m) < 2.5×1010
  Dominated by one lot (T107) – appears to have been assigned the wrong temperature
  High Q0 also for mixed 900/950 °C cavities with pre-treated subassemblies

## Slide 12: Cavity Reprocessing at the Partner Labs
High pressure rinsing to remove field emission or persistent multipacting
  HPR successful on ~60 cavities
    Full analysis ongoing
    Only 3 cavities with field emission remaining.
  Typical success rate ~50% per HPR & retest
    Re-rinsed if any field emission detected, even if above CM admin limit.
  Over 200 HPRs performed – persistence often pays off!
HF acid rinse to remediate low quench field
  Decrease in quench field in some cavities may be caused by oxide growth after many HPRs
  Some cavities also appeared limited by persistent multipacting (especially from CM rebuilds)
  HF rinse has improved Eacc in 18 of 23 cavities (78%)
    FNAL/ANL: HF acid rinse
    JLab: EP fluid flow with no current
  HF rinse has in some cases also removed stubborn field emission

## Slide 13: Cavity Reprocessing at the Cavity Supplier
7 cavities with bellows damage/other issues sent to supplier for re-tanking
  7 qualified after rework
    1 later disqualified due to FE that developed after string assembly process excursion.
6 cavities with performance issues sent to supplier for re-doping
  Remove tank; repeat 50 μm EP, doping, final light EP; re-tank
  Original performance:
    2 with persistent field emission
    4 with low Eacc
  All 6 cavities qualified after rework

## Slide 14: CAVR036 – A Cavity Rework Saga
Feb 2022: Bellows damage detected at supplier
Jun 2022: Replace bellows at supplier
Jul 2022: RF test, FE encountered, HPR
Aug 2022: RF test, FE encountered, HPR
Aug 2022: RF test, low Eacc, disqualified
Early 2024: Re-doping at cavity supplier
April 2024: RF test, qualified
April 2024: Bellows damage detected
Late 2024: Re-tanking at supplier
Feb 2025: RF test, FE encountered, HPR
Mar 2025: RF test, qualified
Late 2025: String assembly process excursion
Dec 2025: RF test, FE encountered, HPR
Feb 2026: RF test, FE encountered, HPR
Feb 2026: RF test, FE encountered
Mar 2026: HF rinse
Mar 2026: RF test, low Q0 and Eacc, suspect issue with HF rinse
Mar 2026: repeat HF rinse
Mar 2026: RF test, Q0 and Eacc recovered, FE encountered
Apr 2026: RF test, FE encountered
Apr 2026: Cavity disqualified
  4 bellows units – 2 doping runs – 12 RF tests – 2 HF rinses – 2 times qualified – 1 attempted string assembly
  Can’t win ‘em all…

## Slide 15: Cryomodule Production at FNAL and JLab
JLab: 10 CMs, 1 rebuilt
FNAL: 13 CMs (+vCM), 1.5 rebuilt

## Slide 16: Cryomodule Rebuilds
After several successful cooldowns during testing, a pickup probe started leaking and vented the beamline
J22: pickup probe leak
A ceramic was chipped on the warm side of the fundamental power coupler. CM string was still in clean room
F26: chipped ceramic
Root cause analysis points to contamination from a nitrogen purge system
F31: excessive FE
All LCLS-II-HE Cryomodule Builds and Rebuilds Complete

## Slide 17: Cavity Performance in Cryomodules – Gradient
Average cavity performance exceeds cryomodule gradient requirement
Most common limiting factor is the administrative 26 MV/m limit!
Usable Gradient
Gradient-Limiting Factor

## Slide 18: Gradient Change from Vertical to Cryomodule Test
Many cavities reached 26.0 MV/m administrative limit
Average ΔEacc not due to field emission or admin limit: -3.1±3.7 MV/m
Eacc Comparison
Change by Category

## Slide 19: Quality Factor Change from Vertical to Cryomodule Test
Minimal Degradation
Average ΔQ0/Q0: -7.6±17.1%
Q0 Comparison
Average Q0 in CM Exceeds Spec
Average Q0 = 2.9±0.6×1010
Q0 in Cryomodule

## Slide 20: Cryomodule Performance
CMs Q-factor Summary
CMs Gradient Summary
(vCM)
(vCM)
Cryomodule Performance Exceeds Spec with a Comfortable Margin

## Slide 21: Cryomodules Received and Stored at SLAC

## Slide 22: High Level Installation Schedule
FY24 SDT
FY25 SDT
LDT
Installation scheduled during Short/Long Downtimes (SDT/LDT)

## Slide 23: Cryomodule Installation Progress

## Slide 24: LCLS-SC: Maintaining High Q0 After 3 Years
Across the linac an average Q0 of 2.8×1010 was observed during commissioning (spec 2.7×1010)
  Q0 is measured cryogenically due to the low Qext in the CM
Low performing CMs can likely be improved by additional CM degaussing
  Degaussing planned as part of LCLS-II-HE installation
Only obvious change in Q0 due to repeated quenches during multipacting processing; returns to normal after fast cooldown cycle
  NB: Q0 is not measured regularly. However, no major changes in heat load observed.
Average CM Q0
From D. Gonnella, SRF2023 MOIAA04

## Slide 25: Loss of Accelerating Gradient Due to Field Emission
CM01 had no FE prior to installation
Notable FE change kicked off campaign of understanding FE changes in the machine
FE data taken with portable GM-detectors that can measure 2 cryomodules at a time
    Reposition during alternate weekly machine accesses, then re-measure
Have lost 99.1 MV of possible amplitude across all cavities
    8.5% of cavities impacted
CM34 is being replaced during this year’s Long Downtime
Chart: Change in cavity amplitude limits
| CM01-2 | CM01-3 | CM01-4 | CM01-5 | CM01-8 | CM03-2 | CM05-4 | CM08-4 | CM10-5 | CM10-6 | CM14-1 | CM18-2 | CM20-8 | CM21-1 | CM21-2 | CM28-6 | CM30-1 | CM30-3 | CM32-8 | CM33-7 | CM34-5 | CM34-6 | CM34-7 | CM34-8 | CM35-8
After 2025 FE measurement | 14.3 | 9.7 | 5 | 12.3 | 16.6 | 9.3 | 10.8 | 16.6 | 7 | 11 | 10.7 | 12.5 | 16.6 | 14 | 11.8 | 16.6 | 10.5 | 16.6 | 16.3 | 8.3 | 12.3 | 9.5 | 8 | 9 | 6.5
After 2024 FE measurement | 1 | 5.7 | 11.6 | 4.3 | 0 | 0 | 0 | 0 | 4.1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1.5 | 0 | 0.3 | 8.3 | 0 | 0 | 0 | 0 | 10.1
Original amplitude | 1.3 | 1.2 | 0 | 0 | 0 | 3.5 | 5.8 | 0 | 5.3 | 5.6 | 1.3 | 3 | 0 | 1 | 2.7 | 0 | 0 | 0 | 0 | 0 | 4.3 | 4.5 | 7.7 | 5 | 0

## Slide 26: Regaining Gradient by Multipacting Processing
  Selected 7 CMs with most promise for gain
  26 out of 48 cavities processed reached the maximum desired amplitude of 21.8 MV
  Performed a fast cooldown after multipacting processing to restore Q0
  Have now run with a new amplitude limit of 18.68 MV (standard limit is 16.6 MV) on multipacting-processed cavities for about a year
  Gradients higher than 16.6 MV will be needed to meet final desired beam energy of 8.0 GeV.
  Extensive multipacting processing is planned for startup during LCLS-II-HE commissioning.
Chart: MP Processing Limitation
| Field emission | Hard quench | Diagnostic saturation
MP Processing Limitation | 2 | 17 | 3
Chart: 
| 26-1 | 26-2 | 26-3 | 26-5 | 26-7 | 16-1 | 16-2 | 16-3 | 16-4 | 16-5 | 16-6 | 16-7 | 16-8 | 29-1 | 29-3 | 29-4 | 29-5 | 29-6 | 29-7 | 29-8 | 20-1 | 20-2 | 20-3 | 20-4 | 20-5 | 20-6 | 20-7 | 24-1 | 24-2 | 24-3 | 24-4 | 24-5 | 24-6 | 24-7 | 24-8 | 22-1 | 22-2 | 22-3 | 22-4 | 22-5 | 22-6 | 22-8 | 27-1 | 27-2 | 27-3 | 27-4 | 27-6 | 27-8
Final Amplitude (MV) | 18.1 | 21.3 | 19.9 | 21.2 | 21 | 19.7 | 21.8 | 18.8 | 18.9 | 21.8 | 21.8 | 21.4 | 21.3 | 21.8 | 21.3 | 21.8 | 21.8 | 21.8 | 20.9 | 21.8 | 21.8 | 21.8 | 20.4 | 21.8 | 21.8 | 21.8 | 21.6 | 20.1 | 21.8 | 21.8 | 21.8 | 21.3 | 21.8 | 21.8 | 21.3 | 21.8 | 21.8 | 21.8 | 21.8 | 18.1 | 17.9 | 21.8 | 21.8 | 21.8 | 21.8 | 21.6 | 18.2 | 18.5

## Slide 27: Static Heat Load Issues – Vacuum Leaks
Suspected leak associated with static heat load of ~300 W on 2 K circuit in L3B upstream section
Sniffer leak detector showed a leak size of 2.1×10-4 Torr-L/sec
Leak was coming from the vent hole in the multilayer bellows at CM21–22. This points to a latent manufacturing defect.
Bellows were previously leak checked at factory and partner lab
Bellows unit has been replaced
Suspected leak associated with static heat load of 100-200 W on 2 K circuit in L3B downstream section
Leaks identified in CM34 and CM35/Endcap. CM34 measured to be significantly (5×) worse
Spraying helium inside CM34 saw peak readings of 1.8×10-7 Torr-L/sec
Leaks are internal to the cryomodules, difficult to isolate, require CM replacement to repair
CM34 is being replaced with spare CM

## Slide 28: CM34 Replacement with F1.3-32
CM34 ready to leave the tunnel
F1.3-32 staged for replacement

## Slide 29: Suspected Air Leaks to Isolation Vacuum
Larger than expected quantity of nitrogen and water detected in isolation vacuum during warmup.
Estimated 3.5 kg of nitrogen ice removed from L3B downstream, the worst section
May have a compounding effect with helium leaks near CM34
Assumed to be an air leak to the isolation vacuum. No conclusive source(s) identified yet.
L3B String
Likely Air
Likely Water

## Slide 30: Beamline Vacuum Leak
Pressure rise in beamline vacuum observed during L2B isolation vacuum venting process
Interpreted as a leak; venting stopped and isolation vacuum pumped out again.
Leak location estimated to be near CM11
Redundant isolation vacuum pumps added as a precaution.
Planned L2B maintenance put on hold; will eventually need to vent the beamline and isolation vacuum in the future to do the maintenance.

## Slide 31: Fixing Tuners and Couplers During LDT
Several cavities turned off or limited in gradient due to tuner and coupler issues
    Some of these limitations are intermittent
12 of 14 cavities will be investigated and repaired if possible during LDT
Cavity | Issue/Malfunction | Repair Work
CM23-8, CM35-7 | Stepper motor malfunctioning | Investigation and repair ongoing; FNAL colleagues visiting SLAC to assist June 1-12
CM25-2 | Limit switches malfunctioning | 
CM28-1 | Cavity tuning issue | 
CM26-2 | Piezo limited range | 
CM26-8 | FPC heating | 
CM32-8, CM33-(3,5,6,7,8) | Stepper motors intermittent | 
CM04-7 | FPC issue | Postponed due to L2B beamline leak
CM05-8 | Tuner issue | 

## Slide 32: Summary
LCLS-II-HE is a major science-driven upgrade to the superconducting linac at SLAC
The SRF cavities for the project were procured from industry with 95% qualification rate
Cavity performance degradation after cryomodule assembly was acceptably small
All 23 cryomodules have been delivered; installation ongoing through end of 2026
Cavity commissioning scheduled to start in March 2027
First light scheduled for October 2027
LCLS-SC has experienced some gradient loss over time due to field emission
Multipacting processing has been used to regain gradient elsewhere; will be done extensively during LCLS-II-HE commissioning
There has not been significant change to the dynamic heat load (Q0) over time
Repairs ongoing to address tuner, coupler, and vacuum leak issues