## Slide 1
Development of 1.3 GHz Cavity for S3FEL project
Yuchao FANG
On behalf of IASF SRFC team
Institute of Advanced Light Source Facilities, Shenzhen
  10th Jun. TTC’2026 Meeting, Paris French

## Slide 2
Outline
01
Introduction
02
Development of 1.3GHz 9cell cavity
04
Summary
03
Testing results

## Slide 3: 1. Introduction

## Slide 4
Guangming District
CHINA
Great Bay Area
S3FEL: Shenzhen Superconducting Soft X-ray Free Electron Laser
Construction unit: Institute of Advanced Light Source Facilities, Shenzhen (IASF)
Construction site: Guangming District, Shenzhen (Guangming science city, cluster of facilities)

## Slide 5
S3FEL: Shenzhen Superconducting Soft X-ray Free Electron Laser
LINAC: Based on SRF TESLA cavities, operating in CW mode Energy: 2.0 - 2.5 GeV LINAC Length: 750 meters S3FEL Total Length: 1700 meters Four Laser beamlines in the first phase Construction period: 7 years
Injector
LINAC
Beamline
Laboratory

## Slide 6
荔湖水库
（原楼村水库）
2017-03-15
2022-08-10
2025-06-26
S3FEL at 2032

## Slide 7
1
S3FEL: Main LINAC
Layout of S3FEL Accelerators
2
SMTF: SC Module Test Facility
3
LTF: LINAC Test Facility

## Slide 8
1.3GHz CMs
5*
1*
18*
S3FEL: Main LINAC
Layout of S3FEL Accelerators
SMTF: SC Module Test Facility
LTF: LINAC Test Facility
1.3GHz cavities
40*
8*
144*

## Slide 9: 2. Development of 1.3GHz 9cell cavity

## Slide 10
Helium Vessel
Embrace Arm
Coupler Port
S3FEL 1.3 GHz 9-cell Dressed Cavity Design
HOM Coupler
2-Phase Pipe
Fast cooling Pipe
Pickup
Technical Specs of 1.3GHz 9 cell Cavity
Parameters | EXFEL | LCLS-Ⅱ | SHINE | DALS | S3FEL
Cavity type | TESLA | TESLA | TESLA | TESLA | TESLA
Operation mode | Pulse | CW | CW | CW | CW
Temperature(K) | 2.0 | 2.0 | 2.0 | 2.0 | 2.0
Lacc(m) | 1.038 | 1.038 | 1.038 | 1.038 | 1.038
f0(MHz) | 1300 | 1300 | 1300 | 1300 | 1300
Q0 | >5E+9（@16MV/m） | 2.7E+10（@16MV/m） | 3.0E+10（@20MV/m） | 2.7E+10（@16MV/m） | 3.0E+10（@20MV/m）
G(Ohm) | 270 | 270 | 270 | 270 | 270
β | 1 | 1 | 1 | 1 | 1
Epk/Eacc | 2 | 2 | 2 | 2 | 2
Bpk/Eacc(mT/MV/m) | 4.26 | 4.26 | 4.26 | 4.26 | 4.26
R/Q(Ohm) | 1036 | 1036 | 1036 | 1036 | 1036
r/Q(Ohm/m) | 998 | 998 | 998 | 998 | 998
Eacc(MV/m) | 23.6 | 16 | 22 | 19 | 19
KCC | 1.87% | 1.87% | 1.87% | 1.87% | 1.87%
LFD() | 1 |  | 1 | 1 | 1
Parameters | EXFEL | LCLS-Ⅱ | SHINE | DALS | S3FEL
Cavity type | TESLA | TESLA | TESLA | TESLA | TESLA
Operation mode | Pulse | CW | CW | CW | CW
Temperature(K) | 2.0 | 2.0 | 2.0 | 2.0 | 2.0
Lacc(m) | 1.038 | 1.038 | 1.038 | 1.038 | 1.038
f0(MHz) | 1300 | 1300 | 1300 | 1300 | 1300
Q0 | >5E+9（@16MV/m） | 2.7E+10（@16MV/m） | 3.0E+10（@20MV/m） | 2.7E+10（@16MV/m） | 3.0E+10（@20MV/m）
G(Ohm) | 270 | 270 | 270 | 270 | 270
β | 1 | 1 | 1 | 1 | 1
Epk/Eacc | 2 | 2 | 2 | 2 | 2
Bpk/Eacc(mT/MV/m) | 4.26 | 4.26 | 4.26 | 4.26 | 4.26
R/Q(Ohm) | 1036 | 1036 | 1036 | 1036 | 1036
r/Q(Ohm/m) | 998 | 998 | 998 | 998 | 998
Eacc(MV/m) | 23.6 | 16 | 22 | 19 | 19
KCC | 1.87% | 1.87% | 1.87% | 1.87% | 1.87%
 | 1 |  | 1 | 1 | 1
Development of 1.3GHz 9cell cavity

## Slide 11
1.3GHz TESLA Superconducting Cavity
三维电场
三维磁场
仿真模型

## Slide 12
Helium vessel assembly
At vender’s factory
Nb material
Provided by vender
Start of purchase bidding
Bare cavity manufacture
At vender’s factory
Post treatment
At vendor’s factory &
With recipe decided by DALS
Bare cavity vertical test
2 Cavities tested at DESY
Full-Jacked cavity vertical test & Horizontal test
At DALS or domestic collaborative Labs
 | Niobium material | Bare cavity | Post treatment | Optical inspection/ Field flatness adjust | Helium vessel welding | Vertical test
In-Factory | Purchased by Factory | Factory | Factory | Factory | Factory | DALS platform
In-Lab | Purchased by S3FEL | Factory | Collaborative platform | S3FEL lab | Factory | DALS platform

## Slide 13
Pretuning system design
Endoscopic system design
Heat treatment furnance
Post-treatment platform layout
EP System
BCP System

## Slide 14
DALS
Vertical & Horizontal Test Platform
OSTEC & OTIC
Bare cavity fabrication
Raw Material Supply
SEWPG--Shanghai 1st Machinery
Bare cavity fabrication
HE Racing(Jinan)
Bare cavity fabrication
HE Racing(Dongguan)
Potential helium vessel welding supplier
S3FEL

## Slide 15: 3.Testing Results

## Slide 16
Factory Acceptance
On-site Acceptance
Vertical Test

## Slide 17
ISO 4 Cleanroom
VT assembly room
VT assembly
VT Dewar
Insert
Monitors
VT control system interface
Cryo-system monitoring
Clean room and Vertical test platform were established at DALS in Dalian City

## Slide 18
Frequency Control
Final frequency：1300.25±0.1MHz

## Slide 19
Vacuum check of cavities before vertical test
 | Pressure before transportation （mbar） | Pressure after transportation （mbar） | Vacuum Leakage (Pa.m3.s-1) | Final pressure （mbar）
CAV 01 | 8.5E-6 (2022.12.05) | 5.1E-4 (2023.04.08) | ≤1.0E-13 | 3.83E-6 (2023.04.08)
CAV04 | 3.00E-6 (2022.12.12) | 2.09E-4 (2023.03.30) | ≤1.0E-13 | 2.37E-6 (2023.03.30)
CAV05 | 8.00E-6 (2022.12.19) | 1.48E-4 (2023.04.10) | ≤1.0E-13 | 3.43E-6 (2023.04.10)
CAV06 | 6.50E-6 (2022.12.19) | 9.46E-5 (2023.04.08) | ≤1.0E-13 | 3.59E-6 (2023.04.08)
CAV02 | 9.1E-7 (2023.01.17) | 8.8E-5 (2023.04.20) | ≤1.0E-13 | 4.24E-6 (2023.04.20)
CAV08 | 6.20E-6 (2023.02.07) | 6.5E-5 (2023.04.20) | ≤1.0E-13 | 4.74E-6 (2023.04.20)
CAV09 | 3.10E-6 (2023.02.15) | 5.5E-5 (2023.04.21) | ≤1.0E-13 | 3.63E-6 (2023.04.21)
CAV10 | 5.0E-6 (2023.02.17) | 1.12E-4 (2023.04.21) | ≤1.0E-13 | 3.37E-6 (2023.04.21)
The vacuum degree decreases by 1~2 orders of magnitude after 100 days of transportation and storage, but the residual vacuum degree is still better than 5.0E-4 mbar；

## Slide 20: Testing Results
Cavity Srl. No. | Cavity Type | Vertical Test Platform
RI0001 | 1.3GHz 9 cell type-A | DALS
RI0003 | 1.3GHz 9 cell type-A | DALS
RI0004 | 1.3GHz 9 cell type-A | DALS
RI0005 | 1.3GHz 9 cell type-A | DALS
RI0006 | 1.3GHz 9 cell type-A | DALS
RI0007 | 1.3GHz 9 cell type-A | DALS
RI0008 | 1.3GHz 9 cell type-A | DALS
RY66 | 1.3GHz 9 cell type-A | DALS
Cavity VT Result for Standard CM

## Slide 21: 4. Summary

## Slide 22
Requirements defined for CW operation based on TESLA-type 1.3 GHz 9-cell cavity;
8 cavities completed: from factory acceptance → on-site inspection → vertical test;
Post-treatment and testing platform established: EP, BCP, pre-tuning, endoscopy, heat treatment etc.
Summary

## Slide 23
THANK YOU!
Welcome to visit Shenzhen！
深圳市光明区新湖街道
圳园路268号A3栋
268 Zhenyuan St, Building A3, Floor 3-6, Guangming District,
Shenzhen, Guangdong, P.R.China

## Slide 24
Engineering Clarity at Extreme Conditions
Bridging Superconducting RF, Cryogenics and System Control
01
Superconducting RF
• SRF Cavities • Power Couplers • Tuners & HOM • Cryomodules • VTS / CMTS
02
Cryogenic Systems
• Helium Refrigeration • Cryogenic Distribution • 2 K & 4.5 K Systems • Cryostats / Cryomodules • Integration & Commissioning
03
System Control
• LLRF Control • Cryogenic Process Control • Machine Protection • Data Acquisition • Diagnostics / Optimization
Applications
Advanced Light Sources • FEL / ERL • Fusion Cryogenics • Industrial Accelerators • Scientific Infrastructure
Superconducting RF + Cryogenic Engineering + System Control = Integrated Turn-Key Solutions
Design • Manufacturing • Integration • Commissioning • Operation Support
SCTech (Shenzhen) Co., Ltd. | Shenzhen, China
Contact: lihan@mail.iasf.ac.cn