## Slide 1
Cavity performance evolution from vertical test to on-line commissioning at SHINE
Jinfang Chen, SARI, CAS
For the SHINE cryomodule team
June 10, 2026
TTC 2026, Gif sur Yvette, France

## Slide 2
Introduction
    Cavity performance in vertical test
    Performance in horizontal test
    Performance in tunnel commissioning
    Influence factors
Summary
Outline

## Slide 3
Cryomodules requirements
  1 CM with one twin-FPC cavity
  1 CM with 8 cavities (FPC-ABBA orientation)
  54 CMs with eight 1.3 GHz high-Q cavities
  2 CMs with eight 3.9 GHz cavities
Introduction
Accelerator design adjustment：
Energy 8.0 GeV，by pass 3.0-4.5 GeV
54 standard 1.3GHz high-Q CMs
20 MV/m, CM average Eacc
100pC
L3:10
L2:18
L1:2
L4:24
Status
Injector commissioned in Q4, 2024
L1 commissioned in Q3-Q4, 2025
L2-L3 installed, under cooling down now

## Slide 4
Updated Requirements for High-Q Cavities and CMs
 | 1.3 GHz high-Q Cavities |  | 1.3 GHz high-Q CMs | 
 | Bef. Oct 2024 | Updated | Bef. Oct 2024 | Updated
Amount | 600 | 530 ordered | 75 | 54 (+2 spares)
Q0 | 2.7E+10 at 16 MV/m | 3.0E+10 at 20 MV/m | 2.7E+10 at 16 MV/m | 3.0E+10 at 20 MV/m
Max Eacc | ≥ 19 MV/m | ≥ 22 MV/m |  | 26 MV/m(adm. limit)
FE onset |  | ≥ 22 MV/m |  | 
Usable voltage |  |  | ≥ 128 MV | ≥ 166 MV

## Slide 5
1.3GHz Dressed Cavity Performance in Vertical Test
Mid-T
<max Eacc>
27.9±2.9 MV/m
Mid-T
< Q0 >
3.3±0.5×1010
（including 0.8 nΩ flange losses）
N-doping
<max Eacc>
25.2±3.5 MV/m
N-doping
< Q0 >
3.4±0.4×1010
（including 0.8 nΩ flange losses）
SHINE updated VT specs
388 dressed cavities tested:
  138 N-doped
  250 mid-T baked
308 dressed cavities qualified (79%)
57 concessionally accepted for CM (15%):
  Eacc≥ 18 MV/m, or Q0≥2.5E+10@20MV/m
Namely, 94% delivered for CM assembly
As of May, 2026

## Slide 6
1.3GHz CM Performance in HT: Gradient and Q-factor
Average Usable Gradient: 24.0 MV/m
Average Usable Voltage: 199.4 MV
Average Q0 at 20 MV/m: 3.3 E+10
36 CMs (except CM10/22/33)：
Admin limit
Note:
CM10-ND, low performance, dissembled and components reused in other CMs
CM22-MT, reworked, under horizontal test; CM33-MT, not tested

## Slide 7
ΔEacc= -1.85 ± 4.00MV/m
Average Q0 at 20 MV/m
Accelerating gradient(CM01-CM39, except CM10,22,33)
Excluding cavities by administrative limit, 26 MV/m since CM04.
Change in Performance from VT to HT
ΔQ0=

## Slide 8
1.3 GHz CM Performance : Field Emission
86 % (247/288) of the cavities showed no degradation in FE performance from VT to HT
  84% FE-Free
  2% (5 cavities) exhibited FE during VT at gradient > 22 MV/m
14% of the cavities experienced FE degradation
  36 newly emerged FE
  5 reduction in FE onset gradient.
  Average FE onset gradient: 18.4 ± 4.7 MV/m
Note
ND: N-Doping
MT: Mid-T baking

## Slide 9
Commissioning for L1 section (1st cooling down)
Cooling down from 300 K to 160 K
June 24, 2025 ~June 30, 2025
160K to 45K (June 30 ~ July 2)
Studies on degaussing and fast cooling
( Nov 15 ~ Dec 28)
Warm up
Dec 29, 2025 ~ Jan 04, 2026
RF and beam commissioning for L1
(July 18, 2025 ~ Nov 14, 2025)
Fast cooling down
(July 02, 2025)
2+1 CMs

## Slide 10
Injector commissioning: 2024 Q4
+ L1 commissioning: 2025 Q3-Q4
i1CM: performance maintained
i8CM: FE onset degradation observed in 3 cavities, including 1 (cav6) further decreased during beam commissioning
Performance change from HT to commissioning: injector CMs
Injector CMs: i1CM + i8CM
Usable gradients not tested to their maximum in tunnel
Courtesy of Yubin Zhao

## Slide 11
Ebeam = 265 MeV
Performance change from HT to commissioning: L1 CMs
Cavities not tested to their maximum Eacc during RF and beam commissioning in tunnel
No FE degradation observed for all the cavities in L1 from HT to tunnel commissioning
(L1-CM02)
(L1-CM03)
L1 two 1.3 GHz CMs need providing:
250 MeV
Two 3.9 GHz CMs need providing:
- 65 MeV

## Slide 12
Contamination? (suspected in 1 CM，CM10)
Vacuum leak (He-inlet pipe in 1 CM, BPM flange in 1 CM, HOM FT in 2 CMs)
HOM heating （typically, > 24 MV/m）
FE degradation (VT  HT, 14% cavities)
Fast Cooling（typical mass flow rate: 65g/s in HT  40 g/s in L1）
Degaussing (under study in tunnel)
Influence factors

## Slide 13
CM10-ND with abnormally low gradients in HT, contaminated?
HT: Q0=3.4E+10 @ 112 MV
Likely contaminated in CM, cause unclear

## Slide 14
4 cavities (RY040-J, RY049-J, RY050-J, RY058-J)，recovered after HPR，reused in CM29 and CM30 ；
4 cavities (RY044-J, RY047-J, RY051-J, RY056-J), still unqualified after HPR, then recovered by light BCP polishing，and installed in CM36.
CM10-ND: cavities have been dissembled, repaired and reused
VT after HPR (8 cav)
VT after recovery by（4 HPR + 4 BCP）
RY058-J
Curve2 rad disappeared
(Flange losses not deducted)
Specs
4 cavities recovered by BCP light polishing：
Average max Eacc:
13.5 23.3 MV/m
~10 MV/m improved

## Slide 15
Degaussing
SHINE Linac tunnel (~ north-south)
HT stands (~ east-west )
HT stand (HTF1, CM31)
Linac tunnel
Degaussing may affect adjacent CMs in tunnel (under study)
（typically, < 1 mGs reading can be achieved after degaussing）
Flux gate location on cav1/5

## Slide 16
High-Q cavities: high average performance achieved on both mid-T baked and N-doped cavities in VT, among them 365 cavities have been accepted for CM assembly so far.
VT HT: 36 CMs have been tested, including 15 ND and 21 MT CMs, reached an average usable gradient at 24.0 MV/m and Q0 of 3.3 E+10 at 20 MV/m. Slightly degradation observed from VT to HT: ~10% in gradient and Q0, ~15% new FE.
HT Tunnel commissioning (not tested to their maximum): No degradation of gradient found in i1CM and 4 CMs in L1 section, however, new FE appeared in 3 cavities of i8CM.
Influence factors: vacuum leak is the main issue caused CM rework, and CM degaussing in tunnel may be challenging due to mutual interference
CMs of L1-L3 are under cooling down, and first lasing is expected within 2026.
Summary

## Slide 17
Thanks for your attention!
Acknowledgments
  Great effort by SHINE cryomodule team members
  Supports from SHINE cryogenic team, general technology team and undulator team
L1
L2
L3