## Slide 1: Recent Results of Nb/Cu Coated Cavities under KEK/CERN Collaboration
TTC Meeting, 11 June 2025
W. Venturini Delsolaro, S. Atieh, K. Brunner, C. P. A. Carlos, M. Garlaschè, C. Hain, C. Pereira, G. J. Rosaz, J. S. Swieszek, A. G. Terricabras (CERN)
H. Araki, A. Yamamoto, M. Yamanaka (KEK)

## Slide 2: KEK/CERN Collaboration on hydroforming for Nb/Cu cavities
  Nb/Cu development since the time of LEP, LHC etc.
  Copper tube hydroforming
  R&D for FCC on HiPIMS Nb coatings
  Bulk Nb cavity development from TRISTAN
  1.3 GHz processing facility
  Bulk Nb hydroforming
  ILC
CERN
KEK

## Slide 3: 1.3 GHz copper cavity fabrication
Measurement
Coating
Hydroforming
Simulation

## Slide 4: Cavities overview
C1, C1b
The first series of full-seamless
KEK14, KEK15
Brazed and welded with CF flanges
C1c, C1d
Hydroforming + Plasma EP
C3, C3b
Mechanical polishing + light EP + Hydroforming

## Slide 5: Results: Q vs E at 1.85 K
Low field (~1 MV/m)
  Q0 ~ 2-5 ×1010 : Comparable to standard bulk Nb cavities (EP + 120℃ baking)
  Exception: KEK14
Mid field (~10 MV/m)
  Q-drops to < 1×1010
  Exception: C3b
Max Eacc
  12-20 MV/m
  Not as high as standard bulk Nb, but comparable to previous Nb/Cu films on bulk machined substrates.

## Slide 6: Attempts to reduce surface roughness
C3 and C3b
  Mechanical polishing and light EP was applied before hydroforming
  No significant improvement on roughness
C1d
  Plasma EP was applied by INFN Legnaro on copper surface after hydroforming
  Surface roughness was reduced
  Macro-structure (orange-peel) remained
C3b
C3
C1d

## Slide 7: Surface roughness
No clear correlation of surface roughness with RF performance on this small data set
 | C1, C1b | C3, C3b | C1d
Treatment | Standard | Mechanical polishing before hydroforming | Plasma EP ~200 um after hydroforming
Surface |  |  | 
Roughness (Ra) | 3.3 μm | 2.7 μm | 0.69 μm
Q0 at 10 MV/m | C1: 3 ×109 C1b: 3 ×109 | C3: 1 ×109 C3b: 3 ×1010 | C1d: 4 ×109

## Slide 8: Origin of Q-drop
C3 and C3b:
  Produced with the same procedure
  Huge difference in RF performance
Q-drop is caused by residual resistance
Slight decrease of T-dependent Rs with Hrf
First signs of anti-Q slope in Nb films?

## Slide 9: What spoiled C3?
C3: one oxidized spot on the copper surface
Black spot appeared on niobium surface
Possible (likely) explanation: local contamination of Nb layer from the substrate defects

## Slide 10: Effect of cool down gradients (thermoelectric currents)
Residual resistance becomes larger in large ΔT cooling condition
Strong increase of Rres with Hrf
ΔT at Tc:
    Spatial temperature difference causes magnetic flux trapping induced by thermoelectric current.

## Slide 11: Rres / ΔT correlation
Large ΔT cooling (without shield)
Small ΔT cooling (with shield)
Small ΔT cooling is essential for bilayer structure cavities.
A simple “shield” is effective for vertical cryostat measurement.

## Slide 12: Transportation, measurement cross checks
C1, C1b and C3b
  Nb coating at CERN
  First RF test at KEK
  Second RF test at CERN
  No degradation after long-distance transportation
  Consistent measurement results

## Slide 13: Summary
  CERN and KEK started a collaboration to explore hydroforming as a possible substrate manufacturing method for Nb/Cu thin film cavities
  The cavities can be fully-seamless: no welding seams on equator, iris and beam-tube.
  Single-cell 1.3 GHz cavities were successfully produced, and RF performances measured both at CERN and KEK.
  In low-field (~1 MV/m), the quality factor is comparable to standard bulk Nb cavities:
  At mid-field (~10 MV/m), most cavities had Q-drop below 1×1010.
  The reason for Q drops are localized defects (same as in previous experience with other manufacturing methods)
  However, one cavity maintained Q0 > 1×1010 up to 17.5 MV/m.
  The surface roughness is degraded by the hydroforming process due to the huge plastic deformations. The well known orange-peel was maintained even after electropolishing and niobium coating.
  However, this has been observed also in the best performing cavity, so its effect on RF performance is unclear.
  The T-dependent Rs is constant or slightly decreases with field, the Q slope comes from the residual resistance
  An important source of residual resistance is magnetic flux trapping caused by thermoelectric current
  In the measurements using a vertical cryostat, this effect can be eliminated by covering the cavity to ensure that cold helium gas does not blow onto only a part of the cavity, and by cooling it slowly.
  The measurement results from CERN and KEK on the same cavities were consistent. Long-distance transport had no effect on the cavity performance.