## Slide 1: Development of a gamma imager for 3D mapping of radioactivity for applications in nuclear industry
PHENIICS FEST 2026
Giulia ROSSO
Thesis director: Dominique TROMSON
Supervisors: Guillaume AMOYAL
Maugan MICHEL
LCIM – Laboratoire Capteurs et Instrumentation pour la Mesure

## Slide 2: Context and objectives
1

## Slide 3
Nuclear decommissioning
Detection of radioactive hotspots through radiation imaging
Planning
Radiological characterization
Decontamination
Dismantling
Preparation for reuse
Final survey
© A. Macario-Barros

## Slide 4
Ionizing radiation imaging
“To see radioactivity”
Detector
Source
Picture
Visible image
2D radioactivity image
Superimposed image
Several advantages in presence of unknown situations:
Visual localization of radioactive hotspots
Quantitative information (dose/activity)
Remote measurements ALARA principle (As Low As Reasonably Achievable)
Pictures from: Euratom project CLEANDEM (grant agreement n° 945335)
No information about depth!
3D imaging

## Slide 5: 3D gamma imaging method
2

## Slide 6
Method’s overview
SOURCE
POV 1
POV 2
POV 3
Recording of 2D gamma images
Move the detector to a different position
Acquisition of images from different points of view (POV)
3D image reprojection
Triangulation
Estimation of the source position
Need to know the detector’s position at each POV!
No prior information about the environment

## Slide 7
Simultaneous Localization And Mapping
SLAM (Simultaneous Localization And Mapping) problem
“Mathematical problem that seeks to determine a robot’s trajectory while simultaneously creating a map of the environment” [1]
Engel et.al. (2016)
[1] A. M. Barros, “Modular device for automated and reliable mapping of indoor installations”.

## Slide 8
SLAM working principle
Input data
Sensor
(LiDAR, visual camera, radar, IMU unit …)
(depth values, color images, acceleration values …)
SLAM algorithm
Trajectory & point cloud estimation
“Pose” (position & rotation)

## Slide 9
Spatial localization
3D gamma imaging and SLAM coupling
Acquisition of multiple gamma images from different points of view
2D gamma imaging processing
3D hotspot back-projection
Position 1
Position 2
Position 3
Radioactive source
Radiological measurement
SLAM algorithm
Detector’s pose
Acquisition of a video of the surrounding scene

## Slide 10
Position 1
Position 2
Position 3
Spatial localization
3D gamma imaging and SLAM coupling
Acquisition of multiple gamma images from different points of view
2D gamma imaging processing
3D hotspot back-projection
Radiological measurement
SLAM algorithm
Acquisition of a video of the surrounding scene
Detector’s pose

## Slide 11: Development of a prototype
3

## Slide 12: Prototype design
MultImodal RAdioactivity imageR for 3D mapping (MIRAR3D)
Dimensions: (18 cm 10 cm 2.8 cm)
Weight = 325 g
Coded-aperture gamma imaging
Timepix3-based hybrid pixel detector
  1 mm CdTe layer
  256 × 256 pixels (55 µm pixel pitch)
  1.56 ns time resolution
4 mm thick MURA rank-7 coded mask tungsten alloy
Simultaneous Localization And Mapping (SLAM)
Intel ® RealSenseTM RGB-D (color and depth) camera
  Evaluation of the imager’s poses (position and rotation) at each measurement point
  3D sparse point cloud of the environment
Picture from: https://kt.cern/technologies/timepix3

## Slide 13: Experimental validation
4

## Slide 14: Experimental set-up
241Am placeholder
133Ba placeholder
25.1 cm
Measurement of 241Am (25.9 MBq) and 133Ba (39.5 MBq) point-like sources
Three POVs measured
25.1 cm distance between sources
241Am
133Ba
94.7 cm
POV 1
POV 2
POV 3

## Slide 15: Sources separation
59 keV
356 keV
81 keV
“spectro-imaging”:
Separation of the two radionuclides by selecting their energy emission in the energy spectrum.
POV 3
POV 2
POV 1
133Ba
241Am
133Ba
241Am
133Ba
241Am
241Am
Eγ=59 keV
133Ba
Eγ,1= 81 keV
Eγ,2= 356 keV

## Slide 16: Sources localization
Source’s volume barycenter
3D spectro-imaging result
Legend:
× POVs
─ Trajectory
∙ Point cloud
POV 3
POV 1
POV 2
241Am volume
133Ba volume
241Am
133Ba
Estimated distance between sources
x
y
z
Absolute error on the distance:
3D gamma spectro-imaging validated!

## Slide 17: Conclusion
5

## Slide 18: Conclusions and perspectives
Development and validation of a 3D gamma imaging method
Development of a dedicated prototype
Experimental validation in laboratory conditions
Tests on spatially extended sources
Correction of angular resolution degradation effects
System integration
Looking for field tests!
Conclusions:
Perspectives:

## Slide 19
Giulia ROSSO
giulia.rosso@cea.fr
Special thanks to: G. AMOYAL, D. FRAS, Q. GENDRE, A. MACARIO BARROS, M. MICHEL, J. PLAGNARD, L. TOSONI, D. TROMSON
Thank you for your attention!

## Slide 20: Hybrid pixel detector calibration
Detector’s working principle
MiniPIX TPX3
Hybrid pixel detector based on CERN Timepix3 ASIC
CdTe
Timepix3 ASIC
pixels
pixel pitch
Illuminated pixel
Detector matrix
Detector matrix
Localized energy deposition:
Broadened energy deposition:
The deposited charge is confined to the pixel of first interaction
Most probable for (photoelectric absorption)
The deposited charge is shared among multiple pixels
Most probable for (Compton scattering)

## Slide 21: Detector energy resolution
Energy (keV) | (%) |  | 
 | Set 1 | Set 2 | Reference
59.5 | 9.3 | 7.3 | 7.6
80.99 | 15.6 | 10.4 | 14.6
661.7 | 4.4 | 3.7 | 3.9
1332.5 | 4.6 | 2.8 | 4.5
Energy (keV) |  |  | 
 | Set 1 | Set 2 | Reference
59.5 | 9.3 | 7.3 | 7.6
80.99 | 15.6 | 10.4 | 14.6
661.7 | 4.4 | 3.7 | 3.9
1332.5 | 4.6 | 2.8 | 4.5

## Slide 22: Coded-aperture imaging
Mask
Detector
Sensitive to ionizing radiation
Sensitive to the position of interaction (pixelized)
Opaque elements (high Z material)
Transparent elements (holes)
Flux 2
Flux 2
Flux 1
Flux 1
Detector
Coded aperture
1.
2.

## Slide 23: Coded-aperture imaging
Mask
Detector
Sensitive to ionizing radiation
Sensitive to the position of interaction (pixelized)
Opaque elements (high Z material)
Transparent elements (holes)
© E. Reynal De Saint-Michel