## Slide 1
Gravure par Plasma du Titane :
une technique innovante pour la miniaturisation des dispositifs médicaux implantables
Rim Ettouri1, Thomas Tillocher1, Philippe Lefaucheux1, Bertrand Boutaud2 et Rémi Dussart1
1 Groupe de Recherches sur l’Energétique des Milieux Ionisés (GREMI), 45067 Orléans, France
2 MISTIC SAS, Issy-les-Moulineaux, 92130, France
rim.ettouri@univ-orleans.fr
MC23 Plasmas industriels pour la microélectronique et les nouveaux matériaux

## Slide 2
Plasma Etching of Titanium :
An Innovative Technique for Miniaturizing Implantable Medical Devices
Rim Ettouri1, Thomas Tillocher1, Philippe Lefaucheux1, Bertrand Boutaud2 et Rémi Dussart1
1 Groupe de Recherches sur l’Energétique des Milieux Ionisés (GREMI), 45067 Orléans, France
2 MISTIC SAS, Issy-les-Moulineaux, 92130, France
rim.ettouri@univ-orleans.fr
MC23 Plasmas industriels pour la microélectronique et les nouveaux matériaux

## Slide 3: Overview
  Introduction to Titanium-based MEMS for Biomedical Applications
  Active Implantable Medical Devices (AIMD)
  Why Titanium?
  Sampling of Current Projects
  Introduction to Microsystems and Microfabrication
  The Microelectronics Paradigm
  Microsystems Fabrication Schemes
  Microsystems Materials Landscape
  Critical Tasks in High-Aspect-Ratio Titanium Deep Etching
  High-Aspect-Ratio (HAR) Titanium Etching
  Study of Etching Residues and Contamination of the Plasma Reactor
  Process Flow Optimization for Reproducible Microfabrication Processes
3
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 4
Introduction to Titanium-based MEMS for Biomedical Applications
4
4
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 5: Active Implantable Medical Devices (AIMD)
5
Rim Ettouri
rim.ettouri@univ-orleans.fr
Devices implanted in the human body > 29 days
Detection, monitoring, treatment of diseases
Implanted as close as possible to organs
Evolution of AIMD
Smaller: Ease implantation by surgeons, closer to areas to be treated
Smarter: more functions (communication, monitoring…)
Longer lifetime
AIMD are embedded in a sealed titanium-based housing
Functionalization of the housing to miniaturize AIMD and integrated more functions

## Slide 6: Why Titanium?
6
The aerospace, chemical and medical sectors rely on titanium’s strength-to-weight ratio, biocompatibility and corrosion properties.
“Bio-MEMS”
Functionalizing the titanium-based housing of medical implants:
Feedthroughs
Embedded capacitive filter
RF Antenna
Electrodes / Sensors
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 7: Why Titanium?
7
“Bio-MEMS”
Functionalizing the titanium-based housing of medical implants:
Feedthroughs
Embedded capacitive filter
RF Antenna
Electrodes / Sensors
Courtesy of MicroPort CRM
(ex SORIN CRM)
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 8: Sampling of Current Projects
8
2021 - 2024
ProTiMe
ARD CVL Microélectronique
Post-process characterization of titanium-based surfaces for implantable medical devices.
C²P²L Platform
Clean room (GREMI)
ProTiMe
Realization and characterization of micro/nano-structured titanium-based surfaces
    Micro/nanostructuring of titanium by etching it through micro/nanostructured polymer surfaces.
    Analysis of the structure and chemical composition of titanium after etching.
Characterization of processed titanium surfaces
    Influence of the nature of the titanium surface on the biocompatibility.
    Study of the contamination the substrate surface.
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 9: Sampling of Current Projects
9
Rim Ettouri
rim.ettouri@univ-orleans.fr
5 µm
Tech2AIM
Turnkey TECHnologies to enable next generation Active Implantable Medical devices
Development of an ultrasound recharging system.
Communication by antenna integrated on the titanium housing.
Electrodes for electrocardiogram (ECG) detection also integrated on the titanium housing.
2020 - 2023

## Slide 10: Overview
  Introduction to Titanium-based MEMS for Biomedical Applications
  Active Implantable Medical Devices (AIMD)
  Why Titanium?
  Sampling of Current Projects
  Introduction to Microsystems and Microfabrication
  The Microelectronics Paradigm
  Microsystems Fabrication Schemes
  Microsystems Materials Landscape
  Critical Tasks in High-Aspect-Ratio Titanium Deep Etching
  High-Aspect-Ratio (HAR) Titanium Etching
  Study of Etching Residues and Contamination of the Plasma Reactor
  Process Flow Optimization for Reproducible Microfabrication Processes
10
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 11
Introduction to Microsystems and Microfabrication
11
11
Courtesy of STMicroelectronics
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 12: The Microelectronics Paradigm
12
Miniaturization
Devices vary from below one micron up to several millimeters.
Trends towards shrinking down size and volumes.
Large Scale Integration
Better shielding of weak electric signals from sensors.
Reduction of connection cables.
Save cost on extra electronics housing.
Batch Fabrication
Lower the price by collective fabrication on planar substrates.
Improved reproducibility, accuracy and reliability.
Courtesy of IBM
Courtesy of Intel
Courtesy of STMicroelectronics
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 13: Microsystems Fabrication Schemes
13
Surface micromachining
Additive
Bottom-up approach
+
Complex multi-layer structures possible
Freedom of integration
-
Constrained to thin layers
Quality of deposited materials can be problematic
Adapted from: Understanding Smart Sensors, Third Edition, 2013, Artech House.
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 14: Microsystems Fabrication Schemes
14
Bulk micromachining
Adapted from: Understanding Smart Sensors, Third Edition, 2013, Artech House.
Subtractive
Top-down approach
+
Greater structural thickness
Good material reliability
-
Constrained to single layer structures
Materials selection is limited
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 15: Microsystems Materials Landscape
 Bulk Titanium .
Single crystalline substrates
Silicon, quartz (crystalline SiO2), gallium arsenide (GaAs), silicon carbide (SiC), lithium niobate (LiNbO3) and sapphire (Al2O3).
Amorphous substrates
Glass, fused silica (pure SiO2; chemically it is identical to quartz) and alumina (Al2O3).
Polymers
SU-8, PDMS, etc.
Metals
Al, Ni, Cu, Au, etc.
15
Silicon & Si-based materials predominate as materials of choice for microsystems, largely due to microelectronics legacy.
 New applications, performance, and cost requirements are driving development of alternatives.
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 16: Overview
  Introduction to Titanium-based MEMS for Biomedical Applications
  Active Implantable Medical Devices (AIMD)
  Why Titanium?
  Sampling of Current Projects
  Introduction to Microsystems and Microfabrication
  The Microelectronics Paradigm
  Microsystems Fabrication Schemes
  Microsystems Materials Landscape
  Critical Tasks in High-Aspect-Ratio Titanium Deep Etching
  High-Aspect-Ratio (HAR) Titanium Etching
  Study of Etching Residues and Contamination of the Plasma Reactor
  Process Flow Optimization for Reproducible Microfabrication Processes
16
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 17
Critical Tasks in High-Aspect-Ratio Titanium Deep Etching
17
17

## Slide 18: Basic Etching Mechanisms
O
ions
F
S
Fx
Ti
F
4
Mask
Substrate
Passivation layer (TiOxFy)
  Synergistic process between chemically active species and energetic ion bombardment.
18
Rim Ettouri
rim.ettouri@univ-orleans.fr
Cl-based plasma chemistry
F-based plasma chemistry

## Slide 19: High-Aspect-Ratio (HAR) Titanium Etching
19
Titanium can be etched provided thermodynamically stable and volatile by-products are formed
Cl-based plasma chemistry
Ti
TiCl
TiCl2
TiCl3
TiCl4
Cl
Cl
Cl
Cl
Endothermic
Ti
TiF
TiF2
TiF3
TiF4
F
F
F
F
F-based plasma chemistry
Low volatility
Each basic scheme presents one limitation
New scheme
Reaction pathway
Synergistic effect between F and Cl radicals*, which increases the etch rate
* d’Agostino et al., Journal of Applied Physics 72 (1992),
TiFxCly
Cl
Cl
35 µm
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 20
(a) SF6:O2:Ar = 80:40:20 (b) SF6:O2:Ar = 80:20:20 (c) SF6:O2:Ar = 60:40:40
Cross sections of 25 μm wide trenches & top view images of 100 μm wide trenches
etched under various O2-based etching environments: (a) SF6:O2:Ar = 80:40:20,
(b) SF6:O2:Ar = 80:20:20 and (c) SF6:O2:Ar = 60:40:40.
Conditions: 10 mTorr, PICP=1200 W, Pbias=250 W and Δt=30 min.
Sequence of colors obtained as a function of oxide layer thickness in the range of 0-90 nm.
[5] Antończak, A.J., Stępak, B., Kozioł, P.E. et al. The influence of process parameters on the laser-induced coloring of titanium. Appl. Phys. A 115, 1003–1013 (2014).
(a) (b) (c)
Effect of Process Gases on Fabricating Anisotropic Ti Structures: Oxidation Behavior of Ti in an SF6/O2/Ar plasma
20
Rim Ettouri
rim.ettouri@univ-orleans.fr
yellow (17nm); light brown (24nm); brown (28nm); violet (31nm); dark blue (40nm); sky blue (46nm); green

## Slide 21
yellow (17nm); light brown (24nm); brown (28nm); violet (31nm); dark blue (40nm); sky blue (46nm); green
(a) SF6:O2:Ar = 80:40:20 (b) SF6:O2:Ar = 80:20:20 (c) SF6:O2:Ar = 60:40:40
Cross sections of 25 μm wide trenches & top view images of 100 μm wide trenches
etched under various O2-based etching environments: (a) SF6:O2:Ar = 80:40:20,
(b) SF6:O2:Ar = 80:20:20 and (c) SF6:O2:Ar = 60:40:40.
Conditions: 10 mTorr, PICP=1200 W, Pbias=250 W and Δt=30 min.
Grazing incidence X-ray diffraction (GiXRD) patterns of a bulk titanium sample after etching in an SF6/O2/Ar = 60:40:40 plasma.
21
Rim Ettouri
rim.ettouri@univ-orleans.fr
Effect of Process Gases on Fabricating Anisotropic Ti Structures: Oxidation Behavior of Ti in an SF6/O2/Ar plasma

## Slide 22: Path to Manufacturability: Influence of the Reactor Walls on the Reproducibility
22
Quartz Reactor Walls
Quartz Shuttle Substrate Holder
Semiconductor processing infrastructure
  Sub-μm resolution & tolerances
  High fabrication yield
  Large scale integration
  Excellent reproducibility
Prevention of contamination and defects
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 23: Path to Manufacturability: Study of Etching Residues and Contamination of the Plasma Reactor
23
(a) (b) (c)
Evolution of Reactor Wall Contamination During Plasma Etching Of Cl2/SF6/O2 after (a) 30, (b) 120 and (c) 360 minutes respectively.
EDS mapping of the composition of the residues deposited in the titanium trench after etching.
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 24: Summary
Fabrication of Titanium-based MEMS for Biomedical Applications
  Titanium: particularly interesting for many applications
  Critical Tasks in High-Aspect-Ratio Titanium Deep Etching
  Anisotropic features can be etched by a fluorine based plasma (SF6/Ar/O2)
    Necessary to heat the substrate for both etching and passivation
    TiOxFy passivation layer at the sidewalls
  High aspect ratio features can be etched by Cl2/SF6/O2 plasma
    Synergistic effect of F and Cl radical
    Increase of both etch rate selectivity to the Ni mask
    Enhancement of anisotropy with addition of O2
    Aspect ratio as high as 12 could be obtained
  Process Flow Optimization for Reproducible Microfabrication Processes
24
Rim Ettouri
rim.ettouri@univ-orleans.fr

## Slide 25: Acknowledgements
  GREMI
  Prof. Rémi Dussart
  Dr. Thomas Tillocher
  Mr. Philippe Lefaucheux
  Dr. Tinghui Zhang
  MISTIC
  Mr. Bertrand Boutaud
  Dr. Gwenaël Bécan
  Mr. Hadrien Philippe
  Ms. Jodie Phung
25
The work has been supported by the Centre val de Loire Region under the PSPC-Regions project n°1, Project Tech2AIM
Turnkey TECHnologies to enable next generation Active Implantable Medical devices
and within the framework of the CERTeM 5.0, through the ProTiMe project.
Rim Ettouri
rim.ettouri@univ-orleans.fr
Thank You!