## Slide 1
Low cost high performances electrochemical analysis
emmanuel.maisonhaute@sorbonne-universite.fr
Project ELABORE (MITI CNRS 2021-2022)
Partners :
Université de Poitiers : Teko Napporn, Dodzi Zigah
IRD : Agnès Aubouy, Alexis Chaigneau, Valérie Grefeuille
LSTE : Martin Aina
Université Abomey-Calavi : Latifou Lagnika
Associations : Physique sans frontières (François Piuzzi), Chimistes sans Frontières (Manuel Cervera-Marzal), Puya Raimondi (Raymond Campagnolo)

## Slide 2
Outline
Résultats sur le terrain
Electrochemical analysis and instrumentation
PassStat 1 : fast and simple
PassStat 2 : low cost and easy to adapt
A few applications in Africa
Some reflexions and perspectives

## Slide 3
What is electrochemistry?
Science and technique of reciprocal transformation of chemical energy and electrical energy.
Dictionnary :
e-
OxS
Oxel
Oxel
RedS
Physical transport
(diffusion, migration, convection)
+ possible chemical steps
Chemical step
Electron transfer
Adsorption?

## Slide 4
Electrochemistry
Métallurgy
Corrosion
Depollution/Envirronnement
(Bio)sensors
Photovoltaic
Electrosynthesis
Nuclear
Fuel cells
Batteries
Biology
Energy
Electronics
Important to train students and implement research in developping countries
Many applications

## Slide 5
Ox
Potential E (V)
Time (s)
Red
Products
k
E0
 = RT/(Fv)
Characteristic timescale:
v : scan rate = slope (en V/s)
Cyclic voltammetry
Exemple:
Initially on Red in solution
Red Ox + e-
Two usual electrochemical techniques

## Slide 6
Cyclic voltammetry
 = RT/(Fv)
Saveant, JM, Elements of Molecular and Biomolecular Electrochemistry
i  D1/2v1/2
Fast system (Nernstian)
Slow electron transfer kinetics
E0
No chemical reaction
E
I
Two usual electrochemical techniques

## Slide 7
Peak to peak difference Ep: 59 mV (à 25°C)
CV : reversible behaviour
Peak for the maximum gradient:
From Amatore, in Organic Electrochemistry, an introduction and a guide
Two usual electrochemical techniques

## Slide 8
Influence of a chemical reaction
No chemical reaction
Fast reaction
Moderately fast reaction
Peak shift
No backward current
Analysis of current peaks provides information on mechanism
Fast scan rates
E0 and k
Red Ox + e-
Products
Slow scan rates
Two usual electrochemical techniques

## Slide 9
Square wave voltammetry
More sensitive !
Two usual electrochemical techniques

## Slide 10
V
A
Ref
CE
W
+5 V
0
-5 V
Générator 1
Générator 2
Potentiometer
Solution
If Iref = 0, Eref = Cte (Nersnt…)
Concept: add a reference electrode for which Iref = 0.
U
I
Control of E = VW – Vref with a voltmeter
Simply adjust U to get the E you want
Then VCE – Vref is no more important
Possible to do I-E curves!
How to properly perform electrochemical experiments: 3 electrode systems

## Slide 11
From funny unconventional teaching…
Le courant passe: electrochemistry@home
https://www.youtube.com/playlist?list=PL4-5RJd2oIUwrvTe3DuxuikS14xa0iJpy
An electrochemical series:
To distract students, colleagues and friends
To suggest students to make experiments at home (in safe conditions)
To do a maximum with a minimum
Raymond Campagnolo (Puya Raymondi)
Why not making a real low cost potentiostat?
As simple as possible!
… to something almost serious.

## Slide 12
A few low cost potentiostats…
Bibliographie
Whitesides
Matsubara
Few descriptions of the limits
Hard to repare
Design complicated
[2] A. Nemiroski, D.C. Christodouleas, J.W. Hennek, A.A. Kumar, E.J. Maxwell, M.T. Fernández-Abedul, G.M. Whitesides, Universal mobile electrochemical detector designed for use in resource-limited applications, PNAS 111 (33) (2014) 11984–11989,
[3] A. Ainla, M.P.S. Mousavi, M.-N. Tsaloglou, J. Redston, J.G. Bell, M.T. Fernández-Abedul, G.M. Whitesides, Open-source potentiostat for wireless electrochemical detection with smartphones, Anal. Chem. 90 (10) (2018) 6240–6246
[4] Y.C. Li, E.L. Melenbrink, G.J. Cordonier, C. Boggs, A. Khan, M.K. Isaac, L.K. Nkhonjera, D. Bahati, S.J. Billinge, S.M. Haile, R.A. Kreuter, R.M. Crable, T.E. Mallouk, An easily fabricated low-cost potentiostat coupled with user-friendly software for introducing students to electrochemical reactions and electroanalytical techniques, J. Chem. Educ. 95 (9) (2018) 1658–1661,
[5] A.A. Rowe, A.J. Bonham, R.J. White, M.P. Zimmer, R.J. Yadgar, T.M. Hobza, J.W. Honea, I. Ben-Yaacov, K.W. Plaxco, M. Wanunu, CheapStat: An opensource, ‘‘do-it-yourself” potentiostat for analytical and educational applications, PLoS ONE 6 (9) (2011) e23783, [6] P. Irving, R. Cecil, M.Z. Yates, MYSTAT: a compact potentiostat/galvanostat for general electrochemistry measurements, HardwareX. 9 (2021) e00163,
[7] C. Mercer, R. Bennett, P. Conghaile, J. Rusling, D. Leech, Glucose biosensor based on open-source wireless microfluidic potentiostat, Sensors Actuators B Chemical.290 (2019) 616–624
[8] G.N. Meloni, Building a microcontroller based potentiostat: a inexpensive and versatile platform for teaching electrochemistry and instrumentation, J. Chem. Educ. 93 (7) (2016) 1320–1322
[9] S.T. Rajendran, E. Scarano, M.H. Bergkamp, A.M. Capria, C.-H. Cheng, K. Sanger, G. Ferrari, L.H. Nielsen, E.-T. Hwu, K. Zór, A. Boisen, Modular, lightweight, wireless potentiostat-on-a-disc for electrochemical detection in centrifugal microfluidics, Anal. Chem. 91 (18) (2019) 11620–11628
[10] M.D.M. Dryden, A.R. Wheeler, D.T. Eddington, DStat: a versatile, open-source potentiostat for electroanalysis and integration, PLoS ONE 10 (10) (2015) e0140349
[11] Y. Matsubara, A small yet complete framework for a potentiostat, galvanostat, and electrochemical impedance spectrometer, J. Chem. Educ. 98 (10) (2021) 3362–3370
[13] https://iorodeo.com/collections/cheapstat-open-source-potentiostat, (n.d.). https://iorodeo.com/collections/cheapstat-open-source-potentiostat.
[14] M.W. Glasscott, M.D. Verber, J.R. Hall, A.D. Pendergast, C.J. McKinney, J.E. Dick, SweepStat: a build-it-yourself, two-electrode potentiostat for
macroelectrode and ultramicroelectrode studies, J. Chem. Educ. 97 (1) (2020) 265–270, https://doi.org/10.1021/acs.jchemed.9b00893.

## Slide 13
Reminder: basic potentiostat scheme
E+
E-
VS
Masse E+ = 0
R6
I
I
0
-VS = RI
-
+
E-
-
+
E+
E-
E
R3
R5
Electrochemical cell
Potential input
Current output
CE
Ref
W
Suiveur
Current/Voltage converter
-E
Simple electronic components
I = 0
3 électrodes scheme, only 3 amplifiers
LMC 6484
R3 = R5
-E

## Slide 14
The most simple !
Necessitates power supplies, generator and oscilloscope
Analog Discovery 2 : 304 €
No dedicated soft for electrochemistry (see however Matsubara)
PassStat 1.0
+
PassStat 1 : fast and simple

## Slide 15
PassStat 1: fast and simple
Performances: PassStat 1.0
8000 Vs-1
Rather fast !
Rather for confirmed electrochemists
1 mM Ferrocene
In CH3CN + 0.1 M TBAPF6
Micrometric electrode
Gabriel Boitel-Aullen
Talk on Thursday !

## Slide 16
Works on computer or through Bluetooth
Price: between 50 and 80 €
Limited scan rate (100 Vs-1)
Limited compliance ( 2.4 V)
Take care of calibration!
V+
V+
GND
TX
TX
RX
RX
A22
A0
GND
V+
Electrodes
Potentiostat
Teensy 3.6 card
Bluetooth HC05
PassStat 2.0, 2.1 et 2.2
PassStat 2: low cost and easy to adapt

## Slide 17
Open source considerations
+
Arduino core
Python interface
PassStat 2: low cost and easy to adapt

## Slide 18
Open source considerations
Application Android conçue avec Android Studio
Paper in Hardware X (Caux et al., 2022)
Dépôt sur Zenono
Wiki sur https://ohwr.org/project/passstat/wikis/home
Projet GitHub
Certifié par l’Open Source Hardware Association
A diffuser sans restrictions !
Stage Anis Achit, 3 mois, étudiant M1 ingénierie
PassStat 2: low cost and easy to adapt

## Slide 19
Validation
OK!!!
Standard conditions:
1 mM Ferrocene
inCH3CN + 0.1 M TBAPF6
Pt  0.5 mm
Fc
Fc+ + 1e-
PassStat 2: low cost and easy to adapt

## Slide 20
Validation: Square Wave Voltammetry
OK!!!
PassStat 2: low cost and esy to adapt
From Compton, understanding electrochemistry
Standard conditions:
1 mM Ferrocene
inCH3CN + 0.1 M TBAPF6
Pt  0.5 mm

## Slide 21
Validation: low currents
Extreme conditions :
Ferrocene 25 µM
in CH3CN + 1 mM TBAPF6
UME Pt  4 µm
Scan rate 20 mVs-1
Very good S/N!
No stray capacitance
Spherical diffusion: q >> r02/D
Insulating
Insulating
Electrode
Planar diffusion: q << r02/D
Electrode
Insulating
Insulating
Spherical diffusion
PassStat 2: low cost and easy to adapt

## Slide 22
A few examples

## Slide 23
Pollutants
Cu, Cd, Pb, As, Hg
Dietary deficiencies
Paludism
Counterfeit drugs
Essential oils
A few examples

## Slide 24
Drugs: paracetamol
Bluetooth OK !
Produits
PBS
PBS
Citric acid
A few examples: paracetamol
Mélicia Caux, L3, 2021

## Slide 25
Anodic stripping of Cu
Results Université d’Abomey-Calavi, october 2021
Internship M1 Florença Wassolua, 2022
A few examples: copper analysis
Cu2+ 1 mM + NaCl 0.100 M
Titration in tap or sea water

## Slide 26
Essential oil: Occimum gratissimum (faux basilic, Bénin)
Collaboration Pr. Latifou Lagnika, UAC and Agnès Aubouy, IRD
A priori contains thymol
Quantification of a complex electrochemical system in a complex matrix
A few examples: plant extract
Calibration with pure thymol

## Slide 27
Djougou
Different thymol composition depending on the area of collection
A few examples: plant extract, essential oil
Calavi
Very different compositions

## Slide 28
Glassy carbon
Pencil lead
A few examples: plant extract, infusion

## Slide 29
Some reflexions…
Commercial systems
PassStat and other open hardwares
Plug and play
Wide range of techniques
Soft
Excellent resolution
Modulables (filters)
Black boxes
Sometimes bulky
Expensive
Not possible to repair yourself
To mount yourself (1h)
Limited functions and soft (to be improved…)
Correct resolution
Low compliance (can be solved)
Easily adaptable at low cost (T, pH, conductivity…)
Totally open source and easy to repair
Light, works on battery (low noise)
Cost-effective for development
and participative science
Perfect to learn instrumentation !

## Slide 30
Some perspectives
Commercial electrodes
Electrode material
Pencil lead
(Activation ?)
Carbon paste
Screen printed electrodes
Modified electrodes
Systems
Essential oils
Infusions
Plants ???
Hard and soft improvements
Integration within other techniques (e.g. spectroscopy)
Metals
UMEs
Spread the approach to other users
Autonomous underwater measurements

## Slide 31
Acknowledgements
Université de Poitiers : Teko Napporn, Dodzi Zigah
IRD : Agnès Aubouy, Alexis Chaigneau, Valérie Grefeuille
LSTE : Martin Aina
Université Abomey-Calavi : Latifou Lagnika
Associations : Physique sans frontières, Chimistes sans Frontières, Puya Raimondi (Raymond Campagnolo)
Etudiants UAC : Solange Imelda AVOSSE, Faridath BOURAIMA, Sonia Jeliba EDA, Aurellle OGOUTEIBO, Candide SINDEDJI, Satar AKADIRI, Jean Haffiz, Cynthia ATTINDEHOU, Mounirou TCHATCHEDRE
Etudiants SU : Mélicia Caux, Anis Achit, Kethsovann Var, Florença Wassolua, Guillaume Nicole, Gabriel Boitel-Aullen