## Slide 1: S001 AGATA capsule scan trapping study vs temperature
G. Duchêne, F. Didierjean, M. Filliger, M. Moukaddam, M-H. Sigward

## Slide 2: Electron trapping in n-type Ge detectors
Energy peak shifts (Mohamad’s talk)
  Detector scanned in vertical position with a 137Cs source (662 keV)
  Gamma-ray spectrum collected at each XY position
  Peak shift plotted vs XY/segment slice
  Anomaly observed in slices 2, 3 and 4 (red arrows)
B. De Canditiis, PhD Univ. Strasbourg, 2020

## Slide 3: Electron trapping in n-type Ge detectors
Normal electron
trapping
Abnormal electron
trapping
B. De Canditiis, PhD
Univ. Strasbourg, 2020

## Slide 4: Electron trapping versus temperature
Trapping versus temperature
  Trap depth = ΔE = EC – ET
  Heat generates vibrations (phonons) which may excite the trapped electrons
  Bolzmann law ~ C*e –(ΔE/2kT)
  Explore in temperature and search temperature at which tail disappears
  ΔE determination
ΔE
EC
EV
ET

## Slide 5
Values for Ge(Li) detectors
Trapping > 5 µs for
ΔE > 0.08 eV
D. Guthknecht, PhD,
Univ. Strasbourg, 1971

## Slide 6: Electron trapping versus temperature
How to do that?
  Screw heat resistances on the cold finger
  Place a PT100 for temperature regulation/temperature monitoring close to the heat resistances
heat
resistances
PT100
cold finger
PT100
Dewar
PT100
crystal

## Slide 7: Electron trapping versus temperature
How to do that?
  Screw heat resistances on the cold finger
  Place a PT100 for temperature regulation/temperature monitoring close to the heat resistances
heat
resistances
PT100
cold finger
PT100
Dewar
PT100
crystal

## Slide 8: Electron trapping versus temperature
How to do that?
  Screw heat resistances on the cold finger
  Place a PT100 for temperature regulation/temperature monitoring close to the heat resistances
heat
resistances
PT100
cold finger

## Slide 9: Electron trapping versus temperature
How to do that?
  Screw heat resistances on the cold finger
  Place a PT100 for temperature regulation/temperature monitoring close to the heat resistances
  Open the preamp chamber
  Connect wires to the resistances and the PT100
heat-
resistance
power wires
Cold finger
PT100
wiring

## Slide 10: Electron trapping versus temperature
How to do that?
  Add heat resistances in the cold finger
  Place a PT100 for temperature regulation/temperature monitoring close to the heat resistances
  Open the preamp chamber
  Connect wires to the resistance and the PT100
  Use the NKR-200 MIRION module for regulated T or a standalone power supply not regulated
  Install a PT100 on the end-cap
  Monitor all temperatures (Michel’s talk)

## Slide 11: Electron trapping versus temperature
MIRION NRK-200 module
Standalone power supply

## Slide 12: Electron trapping versus temperature
MIRION NRK-200 module
Standalone power supply

## Slide 13: Electron trapping versus temperature
T° dét -177,97
T° dét -174,3°
T° dét -171,3°
T° dét -164,5°
T° dét -158,1°
Fixed position with default
at different temperatures
Acquisition 300 sec

## Slide 14: Electron trapping versus temperature
T° dét -177,97
T° dét -174,3°
T° dét -171,3°
T° dét -164,5°
T° dét -158,1°
Fixed position with default
at different temperatures
Acquisition 300 sec

## Slide 15: Electron trapping versus temperature
  Trap occupation at T = 0 K
  Occ = N0
FFD = electron distribution function
g(E) = electron state density in metals
In semi-conductors E is replaced by
EC – EV
In our case E is replaced by ΔE = EC – ET
FFD
g(E)
FFD *g(E)
ET
EC
EV
EV

## Slide 16: Electron trapping versus temperature
  Density of trapped electrons excited to the conduction bande
  FFD *g(E) = C(T)* e –[ΔE/2kT]
  Integrating over all energies in the conduction band we get the number of electrons excited to the conduction band
ΔE
-ΔE/2kT

## Slide 17: Electron trapping versus temperature
  Density of trapped electrons excited to the conduction bande
  FFD *g(E) = C(T)* e –[ΔE/2kT]
  Integrating over all energies in the conduction band we get the number of electrons excited to the conduction band
  Trap with increasing T
  Occ(T) = N0*(1- Ne(T))
  T at N0 /2 (T1/2)
  ΔE = 2kT1/2*ln[4(2π/h²*mekT1/2)3/2]
  Extracted ΔE at T= -164°C
  ΔE = 0.131 eV ~ 130 meV
Very preliminary
To be confirmed
ΔE
-ΔE/2kT
N0
N0/2

## Slide 18: Electron trapping versus temperature
  Density of trapped electrons excited to the conduction bande
  ρ = C(T)* e –[ΔE/2kT]
  Integrating over all energies in the conduction band we get the number of electrons excited to the conduction band
  Trap with increasing T
  Occ(T) = N0 *(1- Ne(T))
  T at mid height (T1/2)
  ΔE = 2kT1/2*ln[4(2π/h²*mekT1/2)3/2]
  Extracted ΔE at T = -164°C
  ΔE = 0.131 eV ~ 130 meV
Very preliminary
To be confirmed
ΔE
-ΔE/2kT
D. Guthknecht, PhD,
Univ. Strasbourg, 1971
ΔE (meV) | Τ (µs)
< 50 | < 0.1
~80 | > 5
160 (gold) | ~106

## Slide 19: Perspectives
Similar study in any point of the scanned zone (same ΔE value everywhere along the default line?)
Check the evolution of normal electron trapping and extraction of the corresponding ΔE value
Compare with existing ΔE values for lattice defaults

## Slide 20: Thank you for attention Thanks to IKP for S001 lending