## Slide 1: X-Ray Characterisation of the TCCS and TCCP Crystals
Eloise Matheson, Santiago Andres Solis Paiva, Mario Di Castro
TWOCRYST Workshop: 12th December 2023

## Slide 2: Contents
X-Ray Characterisation
The Crystals
TCCS crystal results
TCCP crystal results
Future X-Ray Diffractometer

## Slide 3: X-Ray Characterisation of Crystals
Use X-Ray diffractometry to measure;
Bending angle / deflection angle
  Angle between incoming and outgoing planes of a crystal
Miscut angle
  Angle between crystal plane and optical surface
Torsion
  Vertical dependence of crystal plane orientations
  Relies on x-rays having wavelength λ typically on the same order of magnitude as the spacing d between planes in the crystal (1 - 100 angstroms)
  Bragg’s Law: 2 𝑑 sin⁡(θ)=𝑛 λ
Illustration of crystalline diffraction.
d
θ
θ

## Slide 4: X-Ray Characterisation of Crystals
Flat surface
a
a
a
a
a
0
x
Angle
x
Cylindrical surface
0
x
Angle
x
Line plot → cylindrical surface
R , m
s, mrad/mm
s
Diagrams from Yury Gavrikov, ‘Characterization of crystals for LHC’
X’Pert Panalytical machine for X-Ray diffractometer

## Slide 5: X-Ray Characterisation of Crystals
Measurement system necessary improvements
Modernisation of the sample table
  Uncertainty on bending and miscut angles -> ± 2 µrad
  Uncertainty on torsion -> ± 1 µrad
  Thermal stability: < 3 µrad/hour
  Beam parameters 130 µm width, 1150 µm height
Adaption of attachment plate for the crystal holder
Addition of reference mirror
Addition of custom-made digital autocollimator (PNPI)

## Slide 6: X-Ray Characterisation of Crystals
The efficiency of crystals cannot be measured with x-rays, must be tested with beam e.g. SPS-H8 beam
Sources of difference between X-Ray machine and SPS-H8 beam measurements
  Smaller x-ray beam that that uses for SPS-H8, hence more sensitivity to any crystal imperfections as the Bragg’s peak can have higher variations in location
  Holder of the crystal may have undergone stress or changes, physically affecting the parameters (bending or torsion angles)

## Slide 7: The crystals: TCCS and TCCP
TCCS
TCCP
Arrangement
TCCS
TCCP
LHC protons

## Slide 8: TCCS Crystal Results
STF-138 Assembly Sketch (INFN)
Crystal | Miscut (µrad) |  |  | Bending (µrad) |  |  | Torsion (µrad/mm) |  | 
STF-138 | Spec. | INFN | CERN | Spec. | INFN | CERN | Spec. | INFN | CERN
 | < 40 | 1 ± 1 | 2.5 ± 2 | 50 - 55 | 52 ± 2 | 51 ± 2 | < 1 | < 1 | 1.3 ± 1

## Slide 9: TCCP Crystal Results
Crystal | Miscut (µrad) |  |  | Bending (mrad) |  |  | Torsion (µrad/mm) |  | 
2C-3 | Spec. | INFN | CERN | Spec. | INFN | CERN | Spec. | INFN | CERN
 | < 100 | < 200 | N/A | 6 – 7.5 | 6.9 ± 0.3 | 7.07 ± 0.3 | < 200 | < 200 | 12 ± 70
2C-3 Assembly Sketch (INFN)
2C-3 Crystal in Machine

## Slide 10: Future X-Ray Diffractometer
Novel in-house X-Ray based diffractometer under commissioning
Can measure crystal wafers and multi-strip crystals
Planned increased machine efficiency and availability keeping uncertainty levels with ± 1 µrad on the measurements of bending angle, miscut angle and torsion for TCPC crystals
In-house machine layout
Rotational stages and miscut measurement procedure studies
New machine under commissioning for 2024 operations