## Slide 1: Some bulk Cu and aC coating preliminary results obtained with the Vacuum Pilot Sector during 2023 run
Quentin DUONG
TE – VSC – VSM
quentin.duong@cern.ch
06/12/2023
Dr. Vincent BAGLIN (CERN) CERN supervisor
Dr. Gaël SATTONNAY (IJCLab-MAVERICS) Doctoral supervisor
Notes: Hello everybody. I’m Quentin DUONG and it’s been 1 year since I began my contract as a PhD student in the Vacuum Surface and Coating group, studying the ecloud in the Vacuum Pilot Sector. During this presentation I’ll talk about results I got for bulk copper and amorphous coating.

## Slide 2: Outline
Set-up description & physics principles
LHC physics fill’s lifetime
Electron cloud of a hybrid physics fill
Electron energy distribution for bulk Copper
TE – VSC – VSM
Notes: I’ll, therefore, introduce my set-up, give some measurements obtained at 450GeV for 72b and hybrid fills, discuss about a hybrid physics fill during its whole lifetime and I’ll end this presentation by giving a glimpse into the electron energy distribution for copper.

## Slide 3: Disclaimer
I present the observations performed on some of the 2023 fills
Results are therefore preliminary
TE – VSC – VSM
Notes: As a disclaimer, I’d like to stress the point that considering the consequent amount of data I’ve to deal with I can’t give a throrough overview of the whole 2023 run and that the results are therefore preliminary.

## Slide 4: Set-up description & physics principles
TE – VSC – VSM
Notes: Alright, let’s introduce the VPS

## Slide 5: Layout
TE – VSC – VSM
  Vacuum Pilot Sector (VPS)
  Location: close to LHCb
  2 beamlines
  Installed in 2015:
    1.4m-long thick a-C coating ~500nm
    1.4m-long bulk Cu
  a-C
  Cu
  CMS
  ATLAS
  ALICE
  LHCb
  B1
  B2
  VPS
  LHC
Notes: The Vacuum Pilot Sector, or VPS, is a 19m-long setup located in between Q5L8 and Q4L8. It has been initially installed in 2014 but has been modified in 2015, adding the two stations I’ll discuss: a 1.4m-long station of amorphous carbon and one of bulk copper. I’d like to stress the point here that the carbon coating is in my case a thick one of thickness about 500nm.

## Slide 6: Electron cloud build-up
  Electron cloud = Primary + Secondary electrons
  Primary electrons:
    Photoelectrons induced by photoelectric effect
    Free electrons induced by bunch ionising the residual gas
  Secondary electrons: induced by primary electrons, accelerated by successive bunches, hitting the surface
TE – VSC – VSM
Electron
cloud
Image current
Beam pipe
p+
p+
primary e-
photoelectron
e-
secondary e-
ionised gas i+
accelerated e-
free electron e-
PEY
SEY
  Two main material properties for ecloud:
  Photoelectron Yield (PEY)
  Secondary Electron Yield (SEY)
Notes: As a quick reminder of the ecloud build-up process, it begins with pimary electrons consisting in photoelectrons induced by synchrotron radiation hitting the beamscreen, the PEY, and free electrons induced by proton beam ionising the remaining gas. These primary electrons are then accelerated until a certain energy with which they’re likely to rip some more electrons by hitting the surface, the SEY. Those are also accelerated and so on, leading to the ecloud build-up.
Ecloud being the sum of primary and secondary electrons.

## Slide 7: Electron current and energy detection
TE – VSC – VSM
A
A
p+
p+
Electron
cloud
Image current
Beam pipe
[0 ; 350] V
[-1.5 ; 1.5] kV
+9V
  Electron pickups
  Energy Analyser
  Electron pickups
    Electromagnetic shield: 7%-transparency
    Geometrical surface:
    Bulk copper collector
    Positively biased:
    Measurement: « electron cloud current »
  Energy Analyser (only copper station)
    « Retarding Field Analyser (RFA) »
    Electron pickup: with bias
    Biased filtering grid:
    Biased filtering grid:
    Measurement:
    « electron cloud current below »
Cu
Cu
Cu
Cu
Cu
Notes: To fathom some of the characteristics of ecloud has been installed in the setup electron pickups which collect the ecloud current, and electron energy analysers only available for copper station. The pickup consists in a positively biased bulk copper collector placed behind an electromagnetic shield with a 7% transparency. The energy analyser, which is precisely-speaking, a retarding field analyser or RFA consists in an electron pickup in which a biased filtering grid is inserted.

## Slide 8: Total pressure detection
TE – VSC – VSM
p+
p+
Image current
Beam pipe
  Total pressure gauges
    Bayard-Alpert gauges « VGI »
    Measurement: « desorbed gas pressure »
  Total pressure gauges
  Two main material properties for gas increase:
  Photodesorption
  Electrodesorption
ionised gas i+
free electron e-
Electron
cloud
electrodesorbed gas
photodesorbed gas
PSD
ESD
  Pressure elevation during LHC functionning
  Photodesorbed gas: induced by photons hitting the surface
  Electrodesorbed gas: induced by electrons hitting the surface
Notes: In addition to the electron-related instruments, there are also total pressure gauges, Bayard-Alpert ones. They’re aimed at monitoring the pressure rises which can, among other phenomena, be induced by the photodesorption induced by the synchrotron radiation, the PSD, and by the electrodesorption induced by ecloud, the ESD.

## Slide 9: LHC physics fill’s lifetime
TE – VSC – VSM
Notes: Now, let’s have a look on the responses of different filling scheme fills during injection, one consisting in batches of 72 bunches, and one hybrid consisting of a mix of 8b4e and 36 b. These fills were chosen to keep the mean proton per bunch constant.

## Slide 10: Hybrid scheme injection
TE – VSC – VSM
  Time
  Injection phase
1 turn of the LHC: ~89µs
ppb
« hybrid »
7*(8b4e) + 3*(36b)
Beam energy
450
6800
B1 intensity
B2 intensity
Beam energy

## Slide 11: Fill lifetime overview
TE – VSC – VSM
  Time
  A
  B
  C
  D
Beam energy
450
6800
B1 intensity
B2 intensity
Beam energy
2000 | 1.5 | 3 | 0.5
2000 | 1.5 | 3 | 0.5
Orders of magnitude of typical beam parameters
  Phases
  Injection (A): batch injections at 450GeV
  Energy ramp-up (B): beam accelerated until 6800GeV
  Stable beam (C): bunch collisions for the experiments
  Ramp-down (D): beam dumped and RF ramp-down
ppb: proton per bunch

## Slide 12: Electron cloud of a hybrid physics fill
TE – VSC – VSM
Notes: Next, I’ll show a few results obtained for the physics hybrid fill during its whole lifetime.

## Slide 13: Electron cloud response of copper during injection
TE – VSC – VSM
  Electron cloud’s linear behaviour after a certain number of bunches injected
Electron flux
[A.cm-2]
Hybrid trains
11
8
6
5
4
3
2
0
400
800
1200
1600
2000
Number of bunches
1
Notes: Here is plotted the electron flux as a fonction of the number of bunches injected only for copper as carbon signals are too close to the detection limit.
You have the electron flux for 72b in green and on the left side and the hybrid one in blue and on the other side.
Equivalent number of trains are also given on the top and very bottom of the plot.
As you can see, both start to behave linearly after a few hundred of bunches. One advantage of hybrid on 72b is thatit needs more bunches before that linear regime, plus its linear cofficient is also 20 times lower than the other one.

## Slide 14: Photoelectrons during energy ramp-up
TE – VSC – VSM
a-C B1
Cu B1
Electron flux [A.cm-2]
a-C Cu
  Photoelectrons detected by electron pickup at which corresponds to a synchrotron radiation critical energy of 2 eV, matching the work function of copper or carbon (~ 4-5 eV)
Energy
B1 intensity
Total beam intensity [p+]
Beam energy [GeV]
Notes: First, you have here the electron flux versus time plot during to the energy ramp-up. In comparison with previous plots there’s here the beam energy plotted in black with its scale in GeV on the right.
As you can see, electron fluxes curb start to increase around 2500GeV showing the presence of photoelectrons. Indeed, for that beam energy the synchrotron radiation critical energy is about 2eV, which matches the work function of copper or carbon (about 4/5 eV).

## Slide 15: Electron cloud current evolution during a fill
  Multipacting induces orders of magnitude of difference between copper and carbon electron cloud signals.
TE – VSC – VSM
a-C B1
Cu B1
Electron flux [A.cm-2]
B1 intensity
Energy
  Multipacting occuring only for copper
  Bumbed curve for copper
  Flat curve for carbon
  Only photoelectrons at the end of the fill
  Both signals scale with beam intensity
Notes: Now, if the plot is zoomed out to the whole fill lifetime, an indication of multipacting in the copper station can be noticed with this bumped curve while amorphous carbon response stays flat.
Moreover, by taking a look to the very last 30minutes of the fill, one can see a linear relation between electron flux and beam intensity, clueing us in a regime of pure photoelectrons. Moreover, one can also measure a factor 6 between copper photoelectrons flux and carbon one and provided the knowledge of the number of photons absorbed in the stations, one can compute their PEY.

## Slide 16: Electron energy distribution for bulk Copper
TE – VSC – VSM
Hybrid (7*8b4e+5*36b): 9072
Notes: Now, let’s move on the last topic of my talk, the electron energy distribution for copper.

## Slide 17: At Stable Beam: t = 0h
TE – VSC – VSM
TE – VSC – VSM
20/12/2023
Electron energy spectrum: Fill 9072
Normalised signal
  Beam parameters
  Energy:
  Mean bunch intensity:
  Fit results
  Low energy:
  High energy:
Notes: Subsequently, energy scans have been done all along beam lifetime, for instance at the begin of stable beam

## Slide 18: At Stable Beam: t = 6h
TE – VSC – VSM
TE – VSC – VSM
20/12/2023
Electron energy spectrum: Fill 9072
Normalised signal
  Beam parameters
  Energy:
  Mean bunch intensity:
  Fit results
  Low energy:
  High energy:
Notes: After 6h of stable beam

## Slide 19: At Stable Beam: t = 11h
TE – VSC – VSM
TE – VSC – VSM
20/12/2023
Electron energy spectrum: Fill 9072
Normalised signal
  Beam parameters
  Energy:
  Mean bunch intensity:
  Fit results
  Low energy:
  High energy:
Notes: And after 11h

## Slide 20: Impact of the beam intensity at 6800GeV
TE – VSC – VSM
TE – VSC – VSM
20/12/2023
Stable Beam |  |  |  | *
0H | 1.59 | 6 | 203 | 263
6H | 1.23 | 6 | 120 | 166
11H | 0.99 | 6 | 89 | 114
0H | 1.59 | 6 | 203 | 263
6H | 1.23 | 6 | 120 | 166
11H | 0.99 | 6 | 89 | 114
  No impact of the bunch intensity on the energy of the secondary electrons.
  Decrease of the energy gained by accelerated electrons with bunch intensity.
  Kick energy formulae gives a reasonable energy assessment at 30%
*J. Scott Berg. CERN LHC Project Note 97, July 1997.
Notes: For your consideration I’ve just summarised here the different peak energies for the three previous scans and it can be pointed out that the secondary electron energy peak has remained the same while the accelerated electrons energy peak decreased with the proton per bunch.
That is consistent with the formulae given by Scott Berg.

## Slide 21: Summary
TE – VSC – VSM
  Electron cloud current:
  At a certain number of bunches,
  Photoelectrons detected for
  Electron cloud responses for copper orders of magnitude above carbon ones
  Multipacting only in copper
  Electron energy distribution:
  Decrease of the accelerated energy peak with ppb
  Kick energy formula is a reasonable first approximation
  Electrostimulated and Photostimulated Desorption (ESD, PSD) computations with Residual Gas Analyser
  Photoelectron Yield (PY) estimations
  Secondary Electrons Yield (SEY) evaluation with RFA by refining measurements around electron peaks
  Measurements of the electron cloud bunch by bunch with a scope
  Heat load evaluation with calorimeters
  Confrontation of measurements and parameter evaluations against simulations
Perspective

## Slide 22: Backups
TE – VSC – VSM

## Slide 23: Station 2 – Amorphous carbon
TE – VSC – VSM
St 1
St 3
z
Ext
Int
7%
7%
7%
7%
K4
K5
Canon 1-12
K6
Sc3a
K16
K18
K17
Canon 1-12
Sc3b
Liner B23
Liner R25
129
110
124
122
121
120
119
118
116
Dash line : Downer side
Red line : Scope
NU : Unused
cal : calorimeter
xx % : pickup with a xx%-transparent grid
Trig : trigger
ph : photodetector
Blue line : Keithley with variable bias
Position :
X
X
X
X
VGI : Bayard-Alpert Gauge (BAG)
VGRB : Pirani gauge​
VQM : Vacuum Quality Monitor
RGA : Residual Gas Analyser
VGPB : Penning Gauge
N : NEG cartridge
X
VPIA : Vacuum Pump
Unused
VF : Vacuum Flange
.5L8.
B
R
RFA : Retarding Field Analyser
kick : kicker (Electron Kicker Detector)

## Slide 24: Station 3 – New unbaked copper
TE – VSC – VSM
St 2
St 4
z
Ext
Int
7%
7%
7%
50%
15%
50%
7%
15%
Sc1a
K7
Canon 1-12
K8
Sc2a
K26
K9
K27
Canon 1-12
Sc2b
K21
Sc1b
Liner B26
Liner R27
92
73
87
85
84
83
82
81
79
Dash line : Downer side
Red line : Scope
Int
NU : Unused
cal : calorimeter
xx % : pickup with a xx%-transparent grid
Trig : trigger
ph : photodetector
Blue line : Keithley with variable bias
Position :
X
X
X
X
VGI : Bayard-Alpert Gauge (BAG)
VGRB : Pirani gauge​
VQM : Vacuum Quality Monitor
RGA : Residual Gas Analyser
VGPB : Penning Gauge
N : NEG cartridge
X
VPIA : Vacuum Pump
Unused
VF : Vacuum Flange
.5L8.
B
R
RFA : Retarding Field Analyser
kick : kicker (Electron Kicker Detector)

## Slide 25: Station 4 – Old unbaked copper
TE – VSC – VSM
St 3
LHCb
z
Ext
Int
Liner 7
Liner 8
55
36
50
48
46
44
42
7%
0,2%
K23
K24
Canon 1-12
Canon 1-12
SHV3
SHV1
K12 + SHV2
K22 + SHV4
K11
HV BIAS 28
0,2%
7%
Dash line : Downer side
Red line : Scope
Blue line : Keithley with variable bias
NU : Unused
cal : calorimeter
xx % : pickup with a xx%-transparent grid
Trig : trigger
ph : photodetector
4
Position :
X
X
X
X
VGI : Bayard-Alpert Gauge (BAG)
VGRB : Pirani gauge​
VQM : Vacuum Quality Monitor
RGA : Residual Gas Analyser
VGPB : Penning Gauge
N : NEG cartridge
X
VPIA : Vacuum Pump
Unused
VF : Vacuum Flange
.5L8.
B
R
RFA : Retarding Field Analyser
kick : kicker (Electron Kicker Detector)
Pt100 air B
Pt100 air R

## Slide 26: Electron cloud for hybrid fills
TE – VSC – VSM
a-C B1
Cu B1
  Detection limit:
  (Corresponds to the mean baseline plus of its gaussian distribution.)
Electron flux [A.cm-2]
  Copper is 4 orders of magnitude above carbon.
  Carbon below detection limit for 8b4e.
  For copper:
B1 intensity
Total beam intensity [p+]
Notes: As a first result, you’ve here plots of the electron flux in ampere per cm² versus time for both filling schemes and for copper in orange and carbon in grey. Plotted in purple you have the detection limit of my measurements, which is about 10pA and corresponds to the mean baseline of the station plus 1 sigma of its gaussian distribution. Finally you’ll find the beam intensity in blue, withh it scale on the right side.
Two points can be noticed:
One is that copper gives responses orders of magnitude above carbon, which is even under the detection limit for the hybrid fill.
The other is that for copper signals are 20 times higher for 72b than hybrid scheme at same beam intensity.

## Slide 27: Ecloud as a function of beam mean ppb
TE – VSC – VSM
  Multipacting presence
  Barely present for carbon: ecloud mainly dominated by photoelectrons
  Huge contribution for copper
  Assymptotic root method
  Root of linear asymptotic behaviour for high ppb
  Method to assess the variations of multipacting with time
linear
multipacting
regime
primary electrons
regime
a-C B1
new-Cu B1
Electron flux [A.cm-2]
Mean ppb
Mean ppb
  Assymptotic root method is NOT aimed at giving a multipacting threshold as ecloud effect has been identified for lower ppb.
Electron flux vs Mean ppb: Fill 9072

## Slide 28: Electron flux scaling with beam intensity at end of fill
TE – VSC – VSM
Electron flux normalised by beam intensity vs Time: Fill 8741
Normalised electron flux [A.cm-2]
a-C B1
Cu B1

## Slide 29: RFA raw data
TE – VSC – VSM
  RFA scan
  Scanning time:
  Filtering voltage:
  Current at a set voltage :
    current of electrons with an energy under
RFA scan: Fill 9072
Electron current vs Filtering voltage: Fill 9072
Electron current
  The energy spectrum of the electrons can be obtained by derivating this curve.

## Slide 30: During energy ramp-up
TE – VSC – VSM
  No significant impact of the beam energy ramp-up on the electron energy distribution.
 |  |  |  | *
620 | 1.60 | 6.0 | 203 | 269
2674 | 1.60 | 5.8 | 202 | 266
6799 | 1.59 | 5.9 | 203 | 263
620 | 1.60 | 6.0 | 203 | 269
2674 | 1.60 | 5.8 | 202 | 266
6799 | 1.59 | 5.9 | 203 | 263
  Double lognormal fit
  Assess the energy peaks of secondary and accelerated electrons
    Measurement step around : 3-5 eV
    Measurement step around : 30 eV
  Cannot describe measurements in-between peaks and above accelerated one
*J. Scott Berg. CERN LHC Project Note 97, July 1997.
Electron energy spectrum: Fill 9072
Normalised signal
Secondary
electrons
Beam kicked
electrons
Notes: The plot you see on the right gives the energy distribution of the electrons, in blue, and presents two peaks, one around 200eV related to the electons kicked by the beam, and one around 6eV related to secondary electrons.
In order to find the energies related to each peak a double lognormal fit as been used locally around those. However, this fit cannot explains the measurements in-between the peaks and above the second one.
These electron peak energies have been identified at different beam energies and gave similar results, just as for Delta Etot, which is the approximation of the energy gained by electrons in presence of beam established by Scott Berg in its paper. Thus, one can observe no significant impact of the beam energy ramp-up on the electron energy distribution.

## Slide 31: Energy spectrum double lognormal fit
TE – VSC – VSM
Electron energy spectrum: Fill 9072
Normalised signal
  Spectrum properties
  Peak at:
  Low peak: Energy of secondary electrons
  High peak: Energy of accelerated electrons
  Limit of the fit: within [20eV ; 100eV] and above 600eV

## Slide 32: Kick approximation principle
TE – VSC – VSM
  Two different types of motions during the passage of a bunch:
  Autonomous approximation: The electrons are strongly pulled by the bunch and begin to oscillate around it.
  Kick approximation: The electrons haven’t started to move very much before.
  Deciding which approximation to use:
    consists of comparing the oscillation period with respect to the time for the bunch to pass
    is equivalent to introducing a transition radius to delimit the domains of validity of both approximations
  Introduction of the critical radius above which the kick approximation is used, and below which the autonomous approximation is used.
  Mean energy kick within these regions:
  Energy kick for an electron on the wall:
  Total mean energy gain:
Ref: “ Energy Gain in an Electron Cloud During the Passage of a Bunch ” – J. Scott Berg

## Slide 33: Kick approximation formulae
TE – VSC – VSM
  Transverse and longitudinal bunch distributions assumed Gaussian
Beam parameters
  Relativistic beta:
  RMS bunch length:
  RMS beam radius:
  Electron properties
  Radius:
  Rest mass:
Ref: “ Energy Gain in an Electron Cloud During the Passage of a Bunch ” – J. Scott Berg

## Slide 34: Energy ramp-up at 620GeV
TE – VSC – VSM
TE – VSC – VSM
20/12/2023
Electron energy spectrum: Fill 9072
Normalised signal
  Beam parameters
  Energy:
  Intensity:
  Fit results
  Low energy:
  High energy:

## Slide 35: Energy ramp-up at 2074GeV
TE – VSC – VSM
TE – VSC – VSM
20/12/2023
Electron energy spectrum: Fill 9072
Normalised signal
  Beam parameters
  Energy:
  Intensity:
  Fit results
  Low energy:
  High energy:

## Slide 36: Energy ramp-up at 6799GeV
TE – VSC – VSM
TE – VSC – VSM
20/12/2023
Electron energy spectrum: Fill 9072
Normalised signal
  Beam parameters
  Energy:
  Intensity:
  Fit results
  Low energy:
  High energy: